**Check Packing Liquid Level Before Each Use**

**Step-by-Step Process for Setting Packing Level in a Wet-Type Gas Meter**

1. Depressurize gas meter by disconnecting all gas lines.
2. Unscrew filling nozzle cap (Figure 2 - Item 7)
3. Unscrew HPLI Screw Cap, if equipped. (Figure 5 -Item 1)
4. Place nipple into outlet gas connection to depressurize outlet (Figure 6)
5. Drain 500 mL of water into calibrated measuring cup using the drain (Figure 2 - Item 11)
6. Verify meter is level using level on meter (Figure 1)



Figure 1: Meter level


1. Liquid Level Indicator
2. Open ball valve (Figure 2 - Item 10) on rear of gas meter.
3. Using water that was drained into the calibrated measuring cup in step 4, start slowing filling the meter in the filling nozzle (Figure 2 - Item 7)
4. Fill the gas meter housing until the liquid column in the level indictor starts to overflow.
5. Wait for it to stop dripping
6. Close ball valve (Figure 2 - Item 10) on rear of gas meter.
7. Setting Calibration
   1. Read HPLI and mL settings on the calibration sheet.
   2. Using the calibrated measuring cup, fill it with the exact mL amount specified on the calibration sheet.
   3. Start slowing filling the meter in the filling nozzle (Figure 2 - Item 7) with the amount specified in step 6b.
   4. Install HPLI Screw Cap and tighten, if equipped. (Figure 5 -Item 1)
   5. Gently, tilt the meter forward, then place back normally. See (Figure 4)
   6. Verify meter is level using level on meter (Figure 1)
   7. Unscrew HPLI Screw Cap, if equipped. (Figure 5 -Item 1)
   8. Verify HPLI setting listed on calibration sheet. Read the liquid level from the base of the meniscus arc
8. Re-assembly
   1. Install HPLI Screw Cap and tighten, if equipped. (Figure 5 -Item 1)
   2. Install and tighten filling nozzle cap (Figure 2 - Item 7)
   3. Verify Liquid Level Indicator ball valve is closed (Figure 2 - Item 10)
   4. Install Outlet Gas Line
   5. Install Inlet Gas Line
   6. Verify meter is level using level on meter (Figure 1). Keep in mind, the gas lines can alter the level slightly. It must be level when everything is installed.
9. Check for Leaks
   1. Purge system and turn on product. Allow for 5 full rotations on gas meter.
   2. Turn off Water Heater and reset EDU 32 FP Display. Verify display reads 0 Liters
   3. With everything connected (including product), allow it to sit for 1 hour. Verify no leaks.

A blueprint of a circular object with numbers and symbols

Description automatically generated A diagram of a liquid column

Description automatically generated

Figure 2: TG 10 – Backside of Gas Meter Figure 3: Liquid Level Setting

 A measuring device with a ruler

Description automatically generated

Figure 4: Tilt Gas Meter Forward Figure 5: HPLI Screw Cap

A close up of a brass pipe

Description automatically generated

Figure 6: Outlet Nipple Installation