## Test form
Run No. : | [9 empty] | Date:
STORAGE WATER HEATER WITH
INPUT RATINGS ≤75,000
WATER HEATER | [7 empty] | Lab ID no.
Manufacturer |  |  |  | Model No. | [5 empty] | Serial No.
TYPE : | Auto/Storage |  |  |  |  | Input Rating |  |  | 40000 |  |  | BTU/HR
[5 empty] | S T A R T |  |  |  |  | F I N I S H
Ambient Temperature |  |  |  |  | 73.8 |  | ºF |  |  | 74.2 |  | ºF
Relative Humidity |  |  |  |  | 27 |  | % |  |  | 27 |  | %
Gas Pressure (Inlet) |  |  |  |  | 7 |  | in H2O |  |  | 7 |  | in H2O
Corrected Barometric Pressure |  |  |  |  | 29.91 |  | in Hg |  |  | 29.91 |  | in Hg
Corrected Gas Meter Reading |  |  |  |  | 0 |  | cu ft |  |  | 9 |  | cu ft
Total Meter Pressure (Meter) |  |  |  |  | 30.5 |  | in Hg |  |  | 30.5 |  | in Hg
Gas Temperature (Meter) |  |  |  |  | 71 |  | ºF |  |  | 71.5 |  | ºF
Heating Value of Gas |  |  |  |  | 1060 |  | Btu/cu ft
CAPACITY:
Water Heater Weight: |  |  | Empty (Wt) |  | 169 |  |  | Full (Wf) |  | 552 |  |  | pounds
Water Temperature: |  |  |  |  | 70 |  | ºF
Density of water,( based on temperature of | [10 empty] | 8.329 | lbs./ gal.
water during weighing )
Storage Tank Capacity (Vst): |  |  |  |  | 45.9839 |  | gal.
Mean Water Inlet Temperature (Tin) |  |  |  |  | 72 |  | ºF
Mean Hot Water Delivery Temperature (Tdel) |  |  |  |  | 135 |  | ºF
Water Drawn: |  |  | START |  |  |  | FINISH
|  | 0 |  | gal. |  | 10.7 |  |  | gal.
Burner cut in time | [7 empty] | 0 |  | sec
Burner cut out time | [7 empty] | 850 |  | sec
Maximum Mean Tank Temperature Recorded Prior to the Test Draw (To) | [7 empty] | 135 |  | ºF
Maximum Mean Tank Temperature Recorded After Cutout Following the Test Draw (Tmax) | [7 empty] | 138 |  | ºF
CONCENTRATION MEASUREMENTS
|  |  |  | CO2 (%) |  |  |  | NOx (ppm) |  |  |  | CO (ppm)
Last Minute Before Cut-Out |  |  |  | 8.5 |  |  |  | 48 |  |  |  | 0
Next To Last Minute |  |  |  | 8.5 |  |  |  | 48 |  |  |  | 0
Preceding Minute |  |  |  | 8.5 |  |  |  | 48 |  |  |  | 0
Average |  |  |  | 8.5 |  |  |  | 48 |  |  |  | 0
FORM M: CORRECTION OF GAS METER READING AND
DETERMINATION OF FIRING RATE
Corrected Gas Meter Reading |  |  |  | (Start) |  |  |  | 0 |  | cu ft
Corrected Gas Meter Reading |  |  |  | (Finish) |  |  |  | 9 |  | cu ft
Corrected Volume of Gas used | [7 empty] | 9 |  | cu ft
Time of Burner Operation | [7 empty] | 850 |  | sec
Heating Value | [7 empty] | 1060 |  | BTU/cu ft
Firing Rate | [7 empty] | 40404.7059 |  | BTU/HR
FORM K: CALCULATION OF HEAT OUTPUT
Average Of Tdel and Tin | [7 empty] | 103.5 |  | ºF
Volume Of Water Withdrawn | [7 empty] | 10.7 |  | gallons
Density Of Water Withdrawn - Avg. of Tdel and Tin | [7 empty] | 8.282 | #REF! | pounds/ gal
Specific Heat Of Water Withdrawn (Cpi) at Avg. of Tdel & Tin | [7 empty] | 0.99904 | #REF! | BTU/ lbºF
Tank
Storage Tank Capacity (Vst) | [7 empty] | 45.9839 |  | gallons
Average Of Tmax and To | [7 empty] | 136.5 |  | ºF
Density Of Water (Dn) - Avg. of Tmax and To | [7 empty] | 8.215 |  | pounds/ gal
Specific Heat Of Water (Cp2) at Avg. of Tmax & To | [7 empty] | 1.0014 |  | BTU/ lbºF
Heat Output (Ho) | [7 empty] | 6712.4557 |  | BTU
FORM N: CALCULATION OF NOx EMMISSIONS
Average Carbon Number Of Fuel (Cf) | [7 empty] | 1.0671 |  | No. Units
Percent Of CO2 Measured ( C ) | [7 empty] | 8.5 |  | %
NOx Measured ( P ) | [7 empty] | 48 |  | ppm
Corrected Fuel Burned ( F ) | [7 empty] | 9 |  | cu ft
Heat Output ( Ho ) | [7 empty] | 6712.4557 |  | BTU
NOx Emissions | [7 empty] | 42.1012 |  | ng/ J

## Water Density
Temp. | Density | Temp. | cP | cP
Deg. F. | lb./gal. | Deg. F. | (kJ/kgK) | Cal/lb-deg.F
60 | 8.337 | 50 | 4.193 | 1.0026
61 | 8.336 | 59 | 4.186 | 1.001
62 | 8.335 | 68 | 4.182 | 1
63 | 8.335 | 77 | 4.181 | 0.99976
64 | 8.334 | 86 | 4.179 | 0.99928
65 | 8.333 | 95 | 4.178 | 0.99904
66 | 8.332 | 104 | 4.179 | 0.99928
67 | 8.331 | 113 | 4.181 | 0.99976
68 | 8.33 | 122 | 4.182 | 1
69 | 8.329 | 131 | 4.183 | 1.0002
70 | 8.329 | 149 | 4.185 | 1.0007
71 | 8.327 | 158 | 4.188 | 1.0014
72 | 8.326
73 | 8.325
74 | 8.324
75 | 8.323
76 | 8.322
77 | 8.321
78 | 8.319
79 | 8.318
80 | 8.317
81 | 8.316
82 | 8.314
83 | 8.313
84 | 8.312
85 | 8.31
86 | 8.309
87 | 8.307
88 | 8.306
89 | 8.305
90 | 8.303
91 | 8.302
92 | 8.3
93 | 8.299
94 | 8.297
95 | 8.295
96 | 8.294
97 | 8.292
98 | 8.291
99 | 8.289
100 | 8.287
101 | 8.286
102 | 8.284
103 | 8.282
104 | 8.28
105 | 8.278
106 | 8.277
107 | 8.275
108 | 8.273
109 | 8.271
110 | 8.269
111 | 8.268
112 | 8.266
113 | 8.264
114 | 8.262
115 | 8.26
116 | 8.258
117 | 8.256
118 | 8.254
119 | 8.252
120 | 8.25
121 | 8.248
122 | 8.245
123 | 8.243
124 | 8.241
125 | 8.239
126 | 8.237
127 | 8.235
128 | 8.232
129 | 8.23
130 | 8.228
131 | 8.226
132 | 8.224
133 | 8.221
134 | 8.219
135 | 8.217
136 | 8.215
137 | 8.212
138 | 8.21
139 | 8.208
140 | 8.205
141 | 8.203
142 | 8.2
143 | 8.198

## Sheet3
Section
2.0 & 4.1 | Ambient | 65 | 85 | F | (during test cannot vary more than ±7F) | Section 7
2 | Relative humidity | 20 | 65 | %
4.2.1 | Inlet water | 68 | 76 | F |  | Installation
4.3.1 | Inlet water pressure | 40 | 150 | psi
4.4 | Gas pressure | 7 | 10 | in wc
4.4 | HHV | 1015 | 1065 | Btu/cu. ft |  |  | 4" dia vent
[6 empty] | 5ft vent | (largest diameter)
Section 5 | Instrumentation
6 probes - use 24 SUT test probe
8 | Test method
8.1 | Input | -2% | 2%
8.1.3 | Maximum mean tank temp. | 130 | 140 | F
Test procedure
1 | Set unit up with probe, 24" stand pipe and 5 ft vent
2 | Fill tank with water and light pilot
3 | Set thermostat to get maximum mean tank temperature between 130F and 140F after first cut out
4 | Wait for thermostat to satisfy
5 | Draw 10 gallon water at 3 gpm
6 | Wait for thermostat to satisfy
7 | Measure the maximum mean tank temperature every minute till maximum mean tank temperature has reached, To
8 | Record this as time T0, gas meter reading at start of test
9 | Draw 10.7 ± 0.5 gallon @ 3.0 ±0.25
10 | If burner cut in has not occurred when 10.7 gallons has been withdrawn continue withdrawal until cut-in occurs
11 | Record burner cut in time
12 | Record burner cut out time
13 | Record inlet (Tin), outlet (Tdel) water temperatures, gas used, water drawn during the draw period
14 | After cut out wait and record maximum mean tank temperature that was reached, Tmax
Scan interval: 15 secs after draw initiation and then 5 sec intervals
Need following information displayed