## Slide 1
Natural Gas - Calibration
Our ask for bottle gas is below:
Primary Grade Natural Gas – HHV = 1025 BTU/cuft (Verification)
Primary Grade Natural Gas – HHV = 1046 BTU/cuft (Calibration)
Primary Grade Natural Gas – HHV = 1072 BTU/cuft (Verification)
We will calibrate the GC with 1046 gas and verify with the other two gases. We will need to determine a “Standard Operating Procedure” on the number of runs, I know we have it, but maybe place in a folder attached to the GC. A bit more visible.
Each calibration/verification certificate will include all runs at each HHV value
Our typical HHV range in Louisville = 1025-1072. This will give both upper and lower limits and a data point in the middle
You will need to determine the appropriate tank size.

## Slide 2
Natural Gas:
Inficon Natural Gas Calibration ~1046 |  |  | 
Component Name | Percentage | HHV (BTU/cuft) | HHV Contribution
Nitrogen | 0.800% | 0.00 | 0.00
Carbon Dioxide | 0.960% | 0.00 | 0.00
Ethane | 4.000% | 1769.00 | 70.76
Propane | 0.500% | 2516.00 | 12.58
Iso-Butane | 0.200% | 3273.00 | 6.55
Butane | 0.100% | 3268.00 | 3.27
Iso-Pentane | 0.080% | 4009.00 | 3.21
Pentane | 0.080% | 3990.00 | 3.19
Hexane | 0.080% | 4830.00 | 3.86
Heptane | 0.000% | 5550.00 | 0.00
Methane | 93.200% | 1010.00 | 941.32
Total (%) | 100.00% |  | 
Est. Higher Heating Value (BTU/cuft) |  |  | 1044.74
Inficon Natural Gas Verification ~1072 |  |  | 
Component Name | Percentage | HHV (BTU/cuft) | HHV Contribution
Nitrogen | 0.550% | 0.00 | 0.00
Carbon Dioxide | 0.200% | 0.00 | 0.00
Ethane | 6.000% | 1769.00 | 106.14
Propane | 0.500% | 2516.00 | 12.58
Iso-Butane | 0.200% | 3273.00 | 6.55
Butane | 0.200% | 3268.00 | 6.54
Iso-Pentane | 0.050% | 4009.00 | 2.00
Pentane | 0.050% | 3990.00 | 2.00
Hexane | 0.050% | 4830.00 | 2.42
Heptane | 0.000% | 5550.00 | 0.00
Methane | 92.200% | 1010.00 | 931.22
Total (%) | 100.00% |  | 
Est. Higher Heating Value (BTU/cuft) |  |  | 1069.44
Inficon Natural Gas Verification ~1025 |  |  | 
Component Name | Percentage | HHV (BTU/cuft) | HHV Contribution
Nitrogen | 0.000% | 0.00 | 0.00
Carbon Dioxide | 0.400% | 0.00 | 0.00
Ethane | 2.200% | 1769.00 | 38.92
Propane | 0.000% | 2516.00 | 0.00
Iso-Butane | 0.000% | 3273.00 | 0.00
Butane | 0.000% | 3268.00 | 0.00
Iso-Pentane | 0.000% | 4009.00 | 0.00
Pentane | 0.000% | 3990.00 | 0.00
Hexane | 0.000% | 4830.00 | 0.00
Heptane | 0.000% | 5550.00 | 0.00
Methane | 97.400% | 1010.00 | 983.74
Total (%) | 100.00% |  | 
Est. Higher Heating Value (BTU/cuft) |  |  | 1022.66
These represent three different Natural Gas Blends.
The certificate should list the below:
1. Primary Grade +/- 1% Accuracy or better
2. Components Listed on Certificate of Analysis (both component, requested concentration, reported Mole %, and accuracy
3. Expiration Date/Re-calibration Date
4. Measured Heating Value listed on Certificate of Analysis

## Slide 3
Propane Gas - Calibration
Our ask for bottle gas is below:
Primary Grade Propane Gas – HHV = 2510 BTU/cuft (Calibration)
Primary Grade Propane Gas – HHV = 2500 BTU/cuft (Verification)
We will calibrate the GC with 2510 gas and verify with the other gas. We will need to determine a “Standard Operating Procedure” on the number of runs, I know we have it, but maybe place in a folder attached to the GC. A bit more visible.
Each calibration/verification certificate will include all runs at each HHV value
This will give us two data points where different components drop out.
You will need to determine the appropriate tank size.

## Slide 4
Propane
Inficon Propane Gas Calibration ~2510 |  |  |  |  | Current Propane Gas Calibration ~2500 |  |  | 
Component Name | Percentage | HHV (BTU/cuft) | HHV Contribution |  | Component Name | Percentage | HHV (BTU/cuft) | HHV Contribution
Ethane | 2.500% | 1769.00 | 44.23 |  | Ethane | 3.500% | 1769.00 | 61.92
Ethylene | 0.250% | 1633.00 | 4.08 |  | Ethylene | 0.000% | 1633.00 | 0.00
Propane | 93.800% | 2516.00 | 2360.01 |  | Propane | 92.450% | 2516.00 | 2326.04
Propylene | 2.00% | 2390 | 47.80 |  | Propylene | 2.70% | 2390 | 64.53
Iso-Butane | 1.000% | 3273.00 | 32.73 |  | Iso-Butane | 1.100% | 3273.00 | 36.00
Butane | 0.250% | 3268.00 | 8.17 |  | Butane | 0.160% | 3268.00 | 5.23
Iso-Pentane | 0.050% | 4009.00 | 2.00 |  | Iso-Pentane | 0.000% | 4009.00 | 0.00
Pentane | 0.050% | 3990.00 | 2.00 |  | Pentane | 0.000% | 3990.00 | 0.00
iso-Butylene | 0.100% | 3195.00 | 3.20 |  | iso-Butylene | 0.090% | 3195.00 | 2.88
Total (%) | 100.00% |  |  |  | Total (%) | 100.00% |  | 
Est. Higher Heating Value (BTU/cuft) |  |  | 2504.21 |  | Est. Higher Heating Value (BTU/cuft) |  |  | 2496.59
Inficon proposes this blend. I propose we move to this blend for calibration.
We use this blend for verification
(This is the current blend in the Range Lab)
The certificate should list the below:
1. Primary Grade +/- 1% Accuracy or better
2. Components Listed on Certificate of Analysis (both component, requested concentration, reported Mole %, and accuracy
3. Expiration Date/Re-calibration Date
4. Measured Heating Value listed on Certificate of Analysis