## GEA Calibration Data - for AHRI
Equipment Description | Manufacturer | Model Number | Calibration ID | Accuracy | Last Calibrated | Calibration Due
GAS METER/TOTALIZER | Alicat Scientific | MW-40SLPM-D-DB9M-IB | 225125 | ±0.6% of reading | 2021-06-29 | 2022-06-29
GAS TEMPERATURE | Alicat Scientific | MW-40SLPM-D-DB9M-IB | 225125 | ± 1.35 °F | 2021-06-29 | 2022-06-29
GAS PRESSURE | Alicat Scientific | MW-40SLPM-D-DB9M-IB | 225125 | ±0.75% of reading | 2021-06-29 | 2022-06-29
TC1 | YOKOGAWA | MW100-SCPSS-TMQSS Series | AP2/CH 6C-MW100 | ± 0.5 °F | 2021-07-15 | 2022-07-15
TC2 | YOKOGAWA | MW100-SCPSS-TMQSS Series | AP2/CH 6C-MW100 | ± 0.5 °F | 2021-07-15 | 2022-07-15
TC3 | YOKOGAWA | MW100-SCPSS-TMQSS Series | AP2/CH 6C-MW100 | ± 0.5 °F | 2021-07-15 | 2022-07-15
TC4 | YOKOGAWA | MW100-SCPSS-TMQSS Series | AP2/CH 6C-MW100 | ± 0.5 °F | 2021-07-15 | 2022-07-15
TC5 | YOKOGAWA | MW100-SCPSS-TMQSS Series | AP2/CH 6C-MW100 | ± 0.5 °F | 2021-07-15 | 2022-07-15
TC6 | YOKOGAWA | MW100-SCPSS-TMQSS Series | AP2/CH 6C-MW100 | ± 0.5 °F | 2021-07-15 | 2022-07-15
Water Flow Meter | Micro Motion | CMFS040MB67N2BAE2ZZ | 12185089 | ± 0.05 °F | 2021-07-15 | 2022-07-15
Power Analyzer | Yokogawa | WT310E | C2VH05016V | ±(0.1% of reading + 0.05 % of range) | 2021-06-21 | 2022-06-22
TC Data Logger | Agilent Technology | 34972A | MY49008622 | ± 1.8 °F | 2021-06-22 | 2022-06-22
RTD Inlet | JMS Southeast Inc | 3SSDNK11.5BGL318HWRAX | AP2/RTD-6C IN | ± 0.2 °F | 2021-07-14 | 2022-07-14
RTD Outlet | JMS Southeast Inc | 3SSDNK5BGL318HWRAX | AP2/RTD-6C OUT | ± 0.2 °F | 2021-07-14 | 2022-07-14
RTD Ambient | JMS Souttheast Inc | 3SSDNK5BGL318HWRAX | AP2/RTD-6C AMB | ± 0.2 °F | 2021-07-14 | 2022-07-14
Manometer | Dwyer | 477AV | AP2/02N2SC | ±0.5% FS | 2022-02-15 | 2023-02-15
Weight Scale | AVERY WEIGH-TRONIX | ZM303 | AP2/193150481 | 0.1 lb | 2021-08-05 | 2022-08-05
TAPE MEASURE | PITTSBURGH | 12 FEET | AP2/TM3 | 0.125 in | 2021-05-18 | 2022-05-18
Manometer | TPI | 608 | AP2/21740410046 | ±0.5% FS | 2021-12-02 | 2022-12-02

## DOE CFR 430
| DOE Requirements: (Per 10 CFR Appendix E to Subpart B of Part 430) | [5 empty] | Link
Condition to Monitor or Measure: | Description/Details | HPWH | Electric | Gas | Instrumentation |  |  | Electric/HPWH Calibration Requirement (Chamber 5 and 8. 1-4 get as close as possible) |  |  |  |  | Gas WH Calibration Requirement (Chamber 6, & Test stations ED1-4)
[5 empty] | Accuracy | Precision | GEA Req't | Equipment Description | Manufacturer | Model Number | Calibration ID | Accuracy | Equipment Description | Manufacturer | Model Number | Calibration ID | Accuracy
Ambient (Air) Temperature | from Cut-in to Cut-out | 67.5+/-1F Average | N/A | N/A | +/-0.2F | +/-0.1F |  | RTD Ambient | JMS Souttheast Inc | 3SSDNK5BGL318HWRAX | AP2/RTD-6C AMB | ± 0.2 °F | RTD Ambient | JMS Souttheast Inc | 3SSDNK5BGL318HWRAX | AP2/RTD-6B AMB | ± 0.2 °F
TPS (Tenney) chamber service | from Cut-out to Cut-in | 67.5+/-2.5F Average | N/A | N/A
support - ticket system keeps | Continuous through test | 67.5 °F ± 5 °F | 65F to 70F | 65F to 70F
record of issue.
Relative Humidity | Throughout test | 50%+/-5% | N/A | N/A | +/-1.5% |  |  | Chamber Humidity sensor DOES require calibration (yearly). TPS is giving us specs on current sensors |  |  |  |  | N/A | N/A | N/A | N/A | N/A
| After Cut-in and before next Cut-out | 50%+/-2% Average | N/A | N/A |  |  |  | Humidity "TEST" sensor: what model will we use? |  |  |  |  | N/A | N/A | N/A | N/A | N/A
Supply Water Temperature | Throughout test | 58F+/-2F | Same | Same | +/-0.2F | +/-0.1F |  | RTD Inlet | JMS Southeast Inc | 3SSDNK11.5BGL318HWRAX | AP2/RTD-6C IN | ± 0.2 °F | RTD Inlet | JMS Southeast Inc | 3SSDNK11.5BGL318HWRAX | AP2/RTD-6C IN | ± 0.2 °F
Outlet Water Temperature | As described in CFR | 125F+/-5F | Same | Same | +/-0.2F | +/-0.1F |  | RTD Outlet | JMS Southeast Inc | 3SSDNK5BGL318HWRAX | AP2/RTD-6C OUT | ± 0.2 °F | RTD Outlet | JMS Southeast Inc | 3SSDNK5BGL318HWRAX | AP2/RTD-6B OUT | ± 0.2 °F
Supply Water Pressure | Throughout test | 40 to 150 psig | Same | Same | +/-1 psi | +/-0.5 psi |  | N/R |  |  |  |  | N/R
Supply Voltage | 5s after Cut-in till 5s before Cut-out | +/-2% of product rated voltage | Same | N/A |  |  |  | WT310E | Yokogawa | WT310E |  | ± 0.39V | WT310E | Yokogawa | WT310E | AP2/CH 6B-MW100 | ± 0.39V
|  | 120V+/-2.4V | Same | N/A
|  | 240V+/-4.8V | Same | N/A
Natural Gas Supply Pressure | Per Mfgr nameplate. OR | N/A | N/A |  | +/-0.1 in WC | +/-.05 in WC |  | N/A | N/A | N/A | N/A | N/A | GAS PRESSURE | Trafag | ESH 8845 (0-40 inWC) | 2010341 | ±0.04 in W.C
| No specification on nameplate? | N/A | N/A | 7-10 in WC (inches of water column) | [8 empty] | Manometer | Dwyer | 477AV-1 (0-20 in W.C) | AP2/02N2SC | ±0.5% FS
[13 empty] | Manometer | TPI | 608 | AP2/21740410046 | ±0.3 in W.C
Natural Gas Heating Value | At standard temp of 60F, Std Pressure 30 inHg |  |  | ~1025 Btu/SCF | ±1% of the reading |  |  | N/A | N/A | N/A | N/A | N/A | Inficon Fusion Micro GC | Inficon | Fusion Micro GC | 141-402609565-1 | Calibration Gas (% Accuracy)
Propane Gas Supply Pressure | Per Mfgr nameplate. OR | N/A | N/A |  | +/-0.1 in WC | +/-.05 in WC |  | N/A | N/A | N/A | N/A | N/A | GAS PRESSURE | Trafag | ESH 8845 (0-40 inWC) | 2010341 | ±0.04 in W.C
| No specification on nameplate? | N/A | N/A | 11-13 in WC (inches of water column) | [8 empty] | Manometer | Dwyer | 477AV-1 (0-20 in W.C) | AP2/02N2SC | ±0.5% FS
[13 empty] | Manometer | TPI | 608 | AP2/21740410046 | ±0.3 in W.C
Propane Gas Heating Value | At standard temp of 60F, Std Pressure 30 inHg | N/A | N/A | ~2500 Btu/SCF | ±1% of the reading |  |  | N/A | N/A | N/A | N/A | N/A | Inficon Fusion Micro GC | Inficon | Fusion Micro GC | 126-401671322-1 | Calibration Gas (% Accuracy)
Atmospheric Pressure | Record in test report | N/A | N/A |  | +/-0.1 inHg | +/-0.05 inHg |  | N/A | N/A | N/A | N/A | N/A | Heise 901A | Heise | 901A | S9-33177 | ±0.07% of span (20-35 in HG ±0.0105 Inches Hg)
Storage Tank Temperature |  |  |  |  | +/-0.5F | +/-0.25F |  | TC1 - TC10 | YOKOGAWA | MW100-SCPSS-TMQSS Series | AP2/CH 6C-MW100 | ± 0.5 °F | TC1 - TC10 | Yokogawa | MW100-SCPSS-TMQSS Series | AP2/CH 6B-MW100 | ± 0.5 °F
[8 empty] | TC Data Logger | Agilent Technology | 34972A | MY49008622 | ± 0.5 °F | TC Data Logger | Agilent Technology | 34972A | MY49001167 | ± 0.5 °F
Mass Water Flow Rate |  |  |  |  | ±1% of the measured value in mass units per unit time |  |  | Coriolis Flow Meters | Micro Motion | CMFS040MB67N2BAE2ZZ | 12185089 | ±0.05% of rate | Water Flow Meter | Micro Motion | CMFS040MB67N2BAE2ZZ | AP2/12227462 | ±0.05% of rate
Electrical Energy (kWh) |  |  |  | N/A | +/-0.5% of reading |  |  | Power Analyzer | Yokogawa | WT310E | C2VH05016V | ±(0.1% of reading + 0.05 % of range) - Range = 6000Watts | Power Analyzer | Yokogawa | WT310E | C2VK21023V | ±(0.1% of reading + 0.05 % of range) - Range = 750Watts
Gas (Fossil Fuels) Used - CuFt |  | N/A | N/A |  | +/-1% of reading |  |  | N/A | N/A | N/A | N/A | N/A | Ritter TG/10 | Ritter | TG10-1.4571-PVDF | 0.54J.D8B | ± 0.5 %
Mass Measurements | <10 lbs |  |  |  | +/-0.1 lbs
| >10 lbs |  |  |  | +/-0.5% of reading |  |  | Weight Scale | AVERY WEIGH-TRONIX | ZM303 | AP2/193150481 | 0.1 lb | Weight Scale | AVERY WEIGH-TRONIX | ZM303 | AP2/193150481 | 0.1 lb
Gas Temperature |  | N/A | N/A | Not Specified in CFR | [8 empty] | TC1 - TC10 | Yokogawa | MW100-SCPSS-TMQSS Series | AP2/CH 6B-MW100 | ± 0.5 °F
Volume (Tank volumes, etc) |  |  |  |  | +/-2% of volume |  |  | Coriolis Flow Meters | Micro Motion | CMFS040MB67N2BAE2ZZ | 12185089 | ±0.05% of totalized volume | Water Flow Meter | Micro Motion | CMFS040MB67N2BAE2ZZ | AP2/12227462 | ±0.05% of totalized volume
Equipment Description | Manufacturer | Model Number | Calibration ID | Accuracy
GAS METER/TOTALIZER | Alicat Scientific | MW-40SLPM-D-DB9M-IB | 225125 | ±0.6% of reading
GAS TEMPERATURE | Alicat Scientific | MW-40SLPM-D-DB9M-IB | 225125 | ± 1.35 °F
GAS PRESSURE | Alicat Scientific | MW-40SLPM-D-DB9M-IB | 225125 | ±0.75% of reading
TC1 | YOKOGAWA | MW100-SCPSS-TMQSS Series | AP2/CH 6C-MW100 | ± 0.5 °F
TC2 | YOKOGAWA | MW100-SCPSS-TMQSS Series | AP2/CH 6C-MW100 | ± 0.5 °F
TC3 | YOKOGAWA | MW100-SCPSS-TMQSS Series | AP2/CH 6C-MW100 | ± 0.5 °F
TC4 | YOKOGAWA | MW100-SCPSS-TMQSS Series | AP2/CH 6C-MW100 | ± 0.5 °F
TC5 | YOKOGAWA | MW100-SCPSS-TMQSS Series | AP2/CH 6C-MW100 | ± 0.5 °F
TC6 | YOKOGAWA | MW100-SCPSS-TMQSS Series | AP2/CH 6C-MW100 | ± 0.5 °F
Water Flow Meter | Micro Motion | CMFS040MB67N2BAE2ZZ | 12185089 | ± 0.05 °F
Power Analyzer | Yokogawa | WT310E | C2VH05016V | ±(0.1% of reading + 0.05 % of range)
TC Data Logger | Agilent Technology | 34972A | MY49008622 | ± 1.8 °F
RTD Inlet | JMS Southeast Inc | 3SSDNK11.5BGL318HWRAX | AP2/RTD-6C IN | ± 0.2 °F
RTD Outlet | JMS Southeast Inc | 3SSDNK5BGL318HWRAX | AP2/RTD-6C OUT | ± 0.2 °F
RTD Ambient | JMS Souttheast Inc | 3SSDNK5BGL318HWRAX | AP2/RTD-6C AMB | ± 0.2 °F
Manometer | Dwyer | 477AV | AP2/02N2SC | ±0.5% FS
Weight Scale | AVERY WEIGH-TRONIX | ZM303 | AP2/193150481 | 0.1 lb
TAPE MEASURE | PITTSBURGH | 12 FEET | AP2/TM3 | 0.125 in
Manometer | TPI | 608 | AP2/21740410046 | ±0.5% FS
Note on F3: Nelson, Jonathan (GE Appliances, Haier): how close a measurement is to the true (calibration std) value.
Note on G3: Nelson, Jonathan (GE Appliances, Haier): How close repeated measurements of same item are to each other. (repeatability)
Note on C4: Sweatt, Corey (GE Appliances, Haier): Average the full range. No running average
Note on C9: Sweatt, Corey (GE Appliances, Haier): Average the full sample, not a rolling average
Note on M36: Sweatt, Corey (GE Appliances, Haier): If it cannot hit accuracy, need to remove

## ID List
Unique ID Numbers
AP2/RTD-6B AMB
AP2/RTD-6C IN
AP2/RTD-6B OUT
AP2/CH 6B-MW100
2010341
AP2/02N2SC
AP2/21740410046
S9-33177
MY49001167
C2VK21023V
0.54J.D8B
AP2/193150481
AP2/12227462