Check Packing Liquid Level Before Each Use

Step-by-Step Process for Setting Packing Level in a Wet-Type Gas Meter

- Depressurize gas meter by disconnecting all gas lines.

- Unscrew filling nozzle cap (Figure 2 - Item 7)

- Unscrew HPLI Screw Cap, if equipped. (Figure 5 -Item 1)

- Place nipple into outlet gas connection to depressurize outlet (Figure 6)

- Drain 500 mL of water into calibrated measuring cup using the drain (Figure 2 - Item 11)

- Verify meter is level using level on meter (Figure 1)

Figure 1: Meter level

- Liquid Level Indicator

- Open ball valve (Figure 2 - Item 10) on rear of gas meter.

- Using water that was drained into the calibrated measuring cup in step 5, start slowing filling the meter in the filling nozzle (Figure 2 - Item 7)

- Fill the gas meter housing until the liquid column in the level indictor starts to overflow.

- Wait for it to stop dripping

- Close ball valve (Figure 2 - Item 10) on rear of gas meter.

- Setting Calibration

  - Read HPLI and mL settings on the calibration sheet. (Verify total mL value on a scale. 1 mL of water = 1 gram of water)

    - Serial Number: 1.544.7HI

      - HPLI = +0.3

      - Subtract 480 mL or 480 grams

    - Serial Number 1.544.8HI

      - HPLI = +0.5

      - Subtract 290 mL or 290 grams

    - Serial Number 1.544.5HI

      - HPLI = -0.5

      - Subtract 300 mL or 300 grams

    - Serial Number 0.54J.D8B

      - HPLI = +0.0

      - Subtract 540 mL or 540 grams

  - Using the calibrated measuring cup, remove the exact mL amount specified in 8a. Using the drain (Figure 2 - Item 11). If you remove too much, simply place the excess back into the meter.

  - Open ball valve (Figure 2 - Item 10) on rear of gas meter, allow level to stabilize, then close ball valve (Figure 2 - Item 10)

  - Gently, tilt the meter forward, then place back normally. Do not allow fluid to drip out of HPLI. (Figure 4)

  - Verify meter is level using level on meter (Figure 1)

  - Verify HPLI setting listed on calibration sheet. Read the liquid level from the base of the meniscus arc.

  - You may need to add/subtract a maximum of 50 mL to dial into the HPLI setting.

- Re-assembly

  - Install HPLI Screw Cap and tighten, if equipped. (Figure 5 -Item 1)

  - Install and tighten filling nozzle cap (Figure 2 - Item 7)

  - Verify Liquid Level Indicator ball valve is closed (Figure 2 - Item 10)

  - Install Outlet Gas Line

  - Install Inlet Gas Line

  - Verify meter is level using level on meter (Figure 1). Keep in mind, the gas lines can alter the level slightly. It must be level when everything is installed.

- Check for Leaks

  - Check HPLI Screw Cap for leaks (Figure 5 -Item 1)

  - Purge system and turn on product. Allow for 5 full rotations on gas meter.

  - Turn off Water Heater and reset EDU 32 FP Display. Verify display reads 0 Liters

  - With everything connected (including product), allow it to sit for 1 hour. Verify no leaks.

Figure 2: TG 10 – Backside of Gas Meter                                                 Figure 3: Liquid Level Setting

Figure 4: Tilt Gas Meter Forward                                             Figure 5: HPLI Screw Cap

Figure 6: Outlet Nipple Installation