## Sheet1
Nominal (SLPM) | Tol - | Tol +
4 | 3.96 | 4.04
8 | 7.952 | 8.048
12 | 11.928 | 12.072
16 | 15.904 | 16.096
20 | 19.88 | 20.12
24 | 23.856 | 24.144
28 | 27.832 | 28.168
32 | 31.808 | 32.192
36 | 35.784 | 36.216
40 | 39.76 | 40.24
Nominal (inHgA) | Tol - | Tol + | Pressure Accuracy
29.9213 | 29.6969 | 30.1457 | 0.22441
30 | 29.7964 | 30.2036 | 0.2036
31 | 30.7964 | 31.2036 | 0.2036
32 | 31.7964 | 32.2036 | 0.2036
33 | 32.7964 | 33.2036 | 0.2036
34 | 33.7964 | 34.2036 | 0.2036
Nominal (Celcius) | Tol - | Tol + | Temp Accuracy
10 | 9.25 | 10.75 | 0.75
15 | 14.25 | 15.75 | 0.75
20 | 19.25 | 20.75 | 0.75
25 | 24.25 | 25.75 | 0.75
30 | 29.25 | 30.75 | 0.75
SLPM (Nominal) | Nominal (SL) | Tol - | Tol + | Tolerance % Reading
20 | 25 | 24.675 | 25.325 | 1.3%
20 | 50 | 49.35 | 50.65 | 1.3%
20 | 75 | 74.025 | 75.975 | 1.3%
20 | 100 | 98.7 | 101.3 | 1.3%
20 | 125 | 123.375 | 126.625 | 1.3%
20 | 150 | 148.05 | 151.95 | 1.3%
We use inHgA on the meter becasue it does not have in W.C. In general, we run pressures between 3.5-14 in water column.
Our STP conditions are 30 in HG and 60F
See if you can hit the 0.75C spec
I know there is a limit on the tank capacity. What is possible?

## Calculations
Assuming a device whose full scale is 40 SLPM, per Sheet1
Assuming flow is a perfect stair stepping, at a static flow rate for the duration of the time interval
= the final actual uncertainty of the flow rates on the left, as a %RD of the totalizer value
= the upper bound of the uncertainty you'd get if you did ±[1.1% (totalizer reading) + [0.1% (device FS)*(time the totalizer has been running)]]
Second count | Flow reading (SLPM) | uncertainty of flow rate (per device's ±0.6%RD or ±0.1%FS, whichever is greater) | Running totalizer value (SL) | Added uncertainty of totalizer (±0.5%RD) | Uncertainty of all data added during the time interval, 1 second in this case (±SLPM) | Running totalizer uncertainty | Running totalizer uncertainty as a %RD of the running totalizer value (±SLPM) | Assuming 0.6% totalizer reading | Assuming 0.1%FS error only | Added 0.5% Totalizer uncertainty from timing | Upper bound of error | Upper bound as %RD
1 | 20 | 0.12 | 0.33333 | 0.1 | 0.22 | 0.0036667 | 1.1 | 0.002 | 0.00066667 | 0.0016667 | 0.0043333 | 1.3
2 | 20 | 0.12 | 0.66667 | 0.1 | 0.22 | 0.0073333 | 1.1 | 0.004 | 0.0013333 | 0.0033333 | 0.0086667 | 1.3
3 | 20 | 0.12 | 1 | 0.1 | 0.22 | 0.011 | 1.1 | 0.006 | 0.002 | 0.005 | 0.013 | 1.3
4 | 20 | 0.12 | 1.3333 | 0.1 | 0.22 | 0.014667 | 1.1 | 0.008 | 0.0026667 | 0.0066667 | 0.017333 | 1.3
5 | 20 | 0.12 | 1.6667 | 0.1 | 0.22 | 0.018333 | 1.1 | 0.01 | 0.0033333 | 0.0083333 | 0.021667 | 1.3
6 | 20 | 0.12 | 2 | 0.1 | 0.22 | 0.022 | 1.1 | 0.012 | 0.004 | 0.01 | 0.026 | 1.3
7 | 20 | 0.12 | 2.3333 | 0.1 | 0.22 | 0.025667 | 1.1 | 0.014 | 0.0046667 | 0.011667 | 0.030333 | 1.3
8 | 20 | 0.12 | 2.6667 | 0.1 | 0.22 | 0.029333 | 1.1 | 0.016 | 0.0053333 | 0.013333 | 0.034667 | 1.3
9 | 20 | 0.12 | 3 | 0.1 | 0.22 | 0.033 | 1.1 | 0.018 | 0.006 | 0.015 | 0.039 | 1.3
10 | 20 | 0.12 | 3.3333 | 0.1 | 0.22 | 0.036667 | 1.1 | 0.02 | 0.0066667 | 0.016667 | 0.043333 | 1.3
11 | 20 | 0.12 | 3.6667 | 0.1 | 0.22 | 0.040333 | 1.1 | 0.022 | 0.0073333 | 0.018333 | 0.047667 | 1.3
12 | 20 | 0.12 | 4 | 0.1 | 0.22 | 0.044 | 1.1 | 0.024 | 0.008 | 0.02 | 0.052 | 1.3
13 | 20 | 0.12 | 4.3333 | 0.1 | 0.22 | 0.047667 | 1.1 | 0.026 | 0.0086667 | 0.021667 | 0.056333 | 1.3
14 | 20 | 0.12 | 4.6667 | 0.1 | 0.22 | 0.051333 | 1.1 | 0.028 | 0.0093333 | 0.023333 | 0.060667 | 1.3
15 | 20 | 0.12 | 5 | 0.1 | 0.22 | 0.055 | 1.1 | 0.03 | 0.01 | 0.025 | 0.065 | 1.3
16 | 20 | 0.12 | 5.3333 | 0.1 | 0.22 | 0.058667 | 1.1 | 0.032 | 0.010667 | 0.026667 | 0.069333 | 1.3
17 | 20 | 0.12 | 5.6667 | 0.1 | 0.22 | 0.062333 | 1.1 | 0.034 | 0.011333 | 0.028333 | 0.073667 | 1.3
18 | 20 | 0.12 | 6 | 0.1 | 0.22 | 0.066 | 1.1 | 0.036 | 0.012 | 0.03 | 0.078 | 1.3
19 | 20 | 0.12 | 6.3333 | 0.1 | 0.22 | 0.069667 | 1.1 | 0.038 | 0.012667 | 0.031667 | 0.082333 | 1.3
20 | 20 | 0.12 | 6.6667 | 0.1 | 0.22 | 0.073333 | 1.1 | 0.04 | 0.013333 | 0.033333 | 0.086667 | 1.3
21 | 20 | 0.12 | 7 | 0.1 | 0.22 | 0.077 | 1.1 | 0.042 | 0.014 | 0.035 | 0.091 | 1.3
22 | 20 | 0.12 | 7.3333 | 0.1 | 0.22 | 0.080667 | 1.1 | 0.044 | 0.014667 | 0.036667 | 0.095333 | 1.3
23 | 20 | 0.12 | 7.6667 | 0.1 | 0.22 | 0.084333 | 1.1 | 0.046 | 0.015333 | 0.038333 | 0.099667 | 1.3
24 | 20 | 0.12 | 8 | 0.1 | 0.22 | 0.088 | 1.1 | 0.048 | 0.016 | 0.04 | 0.104 | 1.3
25 | 20 | 0.12 | 8.3333 | 0.1 | 0.22 | 0.091667 | 1.1 | 0.05 | 0.016667 | 0.041667 | 0.10833 | 1.3
26 | 20 | 0.12 | 8.6667 | 0.1 | 0.22 | 0.095333 | 1.1 | 0.052 | 0.017333 | 0.043333 | 0.11267 | 1.3
27 | 20 | 0.12 | 9 | 0.1 | 0.22 | 0.099 | 1.1 | 0.054 | 0.018 | 0.045 | 0.117 | 1.3
28 | 20 | 0.12 | 9.3333 | 0.1 | 0.22 | 0.10267 | 1.1 | 0.056 | 0.018667 | 0.046667 | 0.12133 | 1.3
29 | 20 | 0.12 | 9.6667 | 0.1 | 0.22 | 0.10633 | 1.1 | 0.058 | 0.019333 | 0.048333 | 0.12567 | 1.3
30 | 20 | 0.12 | 10 | 0.1 | 0.22 | 0.11 | 1.1 | 0.06 | 0.02 | 0.05 | 0.13 | 1.3
31 | 20 | 0.12 | 10.3333 | 0.1 | 0.22 | 0.11367 | 1.1 | 0.062 | 0.020667 | 0.051667 | 0.13433 | 1.3
32 | 20 | 0.12 | 10.6667 | 0.1 | 0.22 | 0.11733 | 1.1 | 0.064 | 0.021333 | 0.053333 | 0.13867 | 1.3
33 | 20 | 0.12 | 11 | 0.1 | 0.22 | 0.121 | 1.1 | 0.066 | 0.022 | 0.055 | 0.143 | 1.3
34 | 20 | 0.12 | 11.3333 | 0.1 | 0.22 | 0.12467 | 1.1 | 0.068 | 0.022667 | 0.056667 | 0.14733 | 1.3
35 | 20 | 0.12 | 11.6667 | 0.1 | 0.22 | 0.12833 | 1.1 | 0.07 | 0.023333 | 0.058333 | 0.15167 | 1.3
36 | 20 | 0.12 | 12 | 0.1 | 0.22 | 0.132 | 1.1 | 0.072 | 0.024 | 0.06 | 0.156 | 1.3
37 | 20 | 0.12 | 12.3333 | 0.1 | 0.22 | 0.13567 | 1.1 | 0.074 | 0.024667 | 0.061667 | 0.16033 | 1.3
38 | 20 | 0.12 | 12.6667 | 0.1 | 0.22 | 0.13933 | 1.1 | 0.076 | 0.025333 | 0.063333 | 0.16467 | 1.3
39 | 20 | 0.12 | 13 | 0.1 | 0.22 | 0.143 | 1.1 | 0.078 | 0.026 | 0.065 | 0.169 | 1.3
40 | 20 | 0.12 | 13.3333 | 0.1 | 0.22 | 0.14667 | 1.1 | 0.08 | 0.026667 | 0.066667 | 0.17333 | 1.3
41 | 20 | 0.12 | 13.6667 | 0.1 | 0.22 | 0.15033 | 1.1 | 0.082 | 0.027333 | 0.068333 | 0.17767 | 1.3
42 | 20 | 0.12 | 14 | 0.1 | 0.22 | 0.154 | 1.1 | 0.084 | 0.028 | 0.07 | 0.182 | 1.3
43 | 20 | 0.12 | 14.3333 | 0.1 | 0.22 | 0.15767 | 1.1 | 0.086 | 0.028667 | 0.071667 | 0.18633 | 1.3
44 | 20 | 0.12 | 14.6667 | 0.1 | 0.22 | 0.16133 | 1.1 | 0.088 | 0.029333 | 0.073333 | 0.19067 | 1.3
45 | 20 | 0.12 | 15 | 0.1 | 0.22 | 0.165 | 1.1 | 0.09 | 0.03 | 0.075 | 0.195 | 1.3
46 | 20 | 0.12 | 15.3333 | 0.1 | 0.22 | 0.16867 | 1.1 | 0.092 | 0.030667 | 0.076667 | 0.19933 | 1.3
47 | 20 | 0.12 | 15.6667 | 0.1 | 0.22 | 0.17233 | 1.1 | 0.094 | 0.031333 | 0.078333 | 0.20367 | 1.3
48 | 20 | 0.12 | 16 | 0.1 | 0.22 | 0.176 | 1.1 | 0.096 | 0.032 | 0.08 | 0.208 | 1.3
49 | 20 | 0.12 | 16.3333 | 0.1 | 0.22 | 0.17967 | 1.1 | 0.098 | 0.032667 | 0.081667 | 0.21233 | 1.3
50 | 20 | 0.12 | 16.6667 | 0.1 | 0.22 | 0.18333 | 1.1 | 0.1 | 0.033333 | 0.083333 | 0.21667 | 1.3
51 | 20 | 0.12 | 17 | 0.1 | 0.22 | 0.187 | 1.1 | 0.102 | 0.034 | 0.085 | 0.221 | 1.3
52 | 20 | 0.12 | 17.3333 | 0.1 | 0.22 | 0.19067 | 1.1 | 0.104 | 0.034667 | 0.086667 | 0.22533 | 1.3
53 | 20 | 0.12 | 17.6667 | 0.1 | 0.22 | 0.19433 | 1.1 | 0.106 | 0.035333 | 0.088333 | 0.22967 | 1.3
54 | 20 | 0.12 | 18 | 0.1 | 0.22 | 0.198 | 1.1 | 0.108 | 0.036 | 0.09 | 0.234 | 1.3
55 | 20 | 0.12 | 18.3333 | 0.1 | 0.22 | 0.20167 | 1.1 | 0.11 | 0.036667 | 0.091667 | 0.23833 | 1.3
56 | 20 | 0.12 | 18.6667 | 0.1 | 0.22 | 0.20533 | 1.1 | 0.112 | 0.037333 | 0.093333 | 0.24267 | 1.3
57 | 20 | 0.12 | 19 | 0.1 | 0.22 | 0.209 | 1.1 | 0.114 | 0.038 | 0.095 | 0.247 | 1.3
58 | 20 | 0.12 | 19.3333 | 0.1 | 0.22 | 0.21267 | 1.1 | 0.116 | 0.038667 | 0.096667 | 0.25133 | 1.3
59 | 20 | 0.12 | 19.6667 | 0.1 | 0.22 | 0.21633 | 1.1 | 0.118 | 0.039333 | 0.098333 | 0.25567 | 1.3
60 | 20 | 0.12 | 20 | 0.1 | 0.22 | 0.22 | 1.1 | 0.12 | 0.04 | 0.1 | 0.26 | 1.3
61 | 20 | 0.12 | 20.3333 | 0.1 | 0.22 | 0.22367 | 1.1 | 0.122 | 0.040667 | 0.10167 | 0.26433 | 1.3
62 | 20 | 0.12 | 20.6667 | 0.1 | 0.22 | 0.22733 | 1.1 | 0.124 | 0.041333 | 0.10333 | 0.26867 | 1.3
63 | 20 | 0.12 | 21 | 0.1 | 0.22 | 0.231 | 1.1 | 0.126 | 0.042 | 0.105 | 0.273 | 1.3
64 | 20 | 0.12 | 21.3333 | 0.1 | 0.22 | 0.23467 | 1.1 | 0.128 | 0.042667 | 0.10667 | 0.27733 | 1.3
65 | 20 | 0.12 | 21.6667 | 0.1 | 0.22 | 0.23833 | 1.1 | 0.13 | 0.043333 | 0.10833 | 0.28167 | 1.3
66 | 20 | 0.12 | 22 | 0.1 | 0.22 | 0.242 | 1.1 | 0.132 | 0.044 | 0.11 | 0.286 | 1.3
67 | 20 | 0.12 | 22.3333 | 0.1 | 0.22 | 0.24567 | 1.1 | 0.134 | 0.044667 | 0.11167 | 0.29033 | 1.3
68 | 20 | 0.12 | 22.6667 | 0.1 | 0.22 | 0.24933 | 1.1 | 0.136 | 0.045333 | 0.11333 | 0.29467 | 1.3
69 | 20 | 0.12 | 23 | 0.1 | 0.22 | 0.253 | 1.1 | 0.138 | 0.046 | 0.115 | 0.299 | 1.3
70 | 20 | 0.12 | 23.3333 | 0.1 | 0.22 | 0.25667 | 1.1 | 0.14 | 0.046667 | 0.11667 | 0.30333 | 1.3
71 | 20 | 0.12 | 23.6667 | 0.1 | 0.22 | 0.26033 | 1.1 | 0.142 | 0.047333 | 0.11833 | 0.30767 | 1.3
72 | 20 | 0.12 | 24 | 0.1 | 0.22 | 0.264 | 1.1 | 0.144 | 0.048 | 0.12 | 0.312 | 1.3
73 | 20 | 0.12 | 24.3333 | 0.1 | 0.22 | 0.26767 | 1.1 | 0.146 | 0.048667 | 0.12167 | 0.31633 | 1.3
74 | 20 | 0.12 | 24.6667 | 0.1 | 0.22 | 0.27133 | 1.1 | 0.148 | 0.049333 | 0.12333 | 0.32067 | 1.3
75 | 20 | 0.12 | 25 | 0.1 | 0.22 | 0.275 | 1.1 | 0.15 | 0.05 | 0.125 | 0.325 | 1.3
76 | 20 | 0.12 | 25.3333 | 0.1 | 0.22 | 0.27867 | 1.1 | 0.152 | 0.050667 | 0.12667 | 0.32933 | 1.3
77 | 20 | 0.12 | 25.6667 | 0.1 | 0.22 | 0.28233 | 1.1 | 0.154 | 0.051333 | 0.12833 | 0.33367 | 1.3
78 | 20 | 0.12 | 26 | 0.1 | 0.22 | 0.286 | 1.1 | 0.156 | 0.052 | 0.13 | 0.338 | 1.3
79 | 20 | 0.12 | 26.3333 | 0.1 | 0.22 | 0.28967 | 1.1 | 0.158 | 0.052667 | 0.13167 | 0.34233 | 1.3
80 | 20 | 0.12 | 26.6667 | 0.1 | 0.22 | 0.29333 | 1.1 | 0.16 | 0.053333 | 0.13333 | 0.34667 | 1.3
81 | 20 | 0.12 | 27 | 0.1 | 0.22 | 0.297 | 1.1 | 0.162 | 0.054 | 0.135 | 0.351 | 1.3
82 | 20 | 0.12 | 27.3333 | 0.1 | 0.22 | 0.30067 | 1.1 | 0.164 | 0.054667 | 0.13667 | 0.35533 | 1.3
83 | 20 | 0.12 | 27.6667 | 0.1 | 0.22 | 0.30433 | 1.1 | 0.166 | 0.055333 | 0.13833 | 0.35967 | 1.3
84 | 20 | 0.12 | 28 | 0.1 | 0.22 | 0.308 | 1.1 | 0.168 | 0.056 | 0.14 | 0.364 | 1.3
85 | 20 | 0.12 | 28.3333 | 0.1 | 0.22 | 0.31167 | 1.1 | 0.17 | 0.056667 | 0.14167 | 0.36833 | 1.3
86 | 20 | 0.12 | 28.6667 | 0.1 | 0.22 | 0.31533 | 1.1 | 0.172 | 0.057333 | 0.14333 | 0.37267 | 1.3
87 | 20 | 0.12 | 29 | 0.1 | 0.22 | 0.319 | 1.1 | 0.174 | 0.058 | 0.145 | 0.377 | 1.3
88 | 20 | 0.12 | 29.3333 | 0.1 | 0.22 | 0.32267 | 1.1 | 0.176 | 0.058667 | 0.14667 | 0.38133 | 1.3
89 | 20 | 0.12 | 29.6667 | 0.1 | 0.22 | 0.32633 | 1.1 | 0.178 | 0.059333 | 0.14833 | 0.38567 | 1.3
90 | 20 | 0.12 | 30 | 0.1 | 0.22 | 0.33 | 1.1 | 0.18 | 0.06 | 0.15 | 0.39 | 1.3
91 | 20 | 0.12 | 30.3333 | 0.1 | 0.22 | 0.33367 | 1.1 | 0.182 | 0.060667 | 0.15167 | 0.39433 | 1.3
92 | 20 | 0.12 | 30.6667 | 0.1 | 0.22 | 0.33733 | 1.1 | 0.184 | 0.061333 | 0.15333 | 0.39867 | 1.3
93 | 20 | 0.12 | 31 | 0.1 | 0.22 | 0.341 | 1.1 | 0.186 | 0.062 | 0.155 | 0.403 | 1.3
94 | 20 | 0.12 | 31.3333 | 0.1 | 0.22 | 0.34467 | 1.1 | 0.188 | 0.062667 | 0.15667 | 0.40733 | 1.3
95 | 20 | 0.12 | 31.6667 | 0.1 | 0.22 | 0.34833 | 1.1 | 0.19 | 0.063333 | 0.15833 | 0.41167 | 1.3
96 | 20 | 0.12 | 32 | 0.1 | 0.22 | 0.352 | 1.1 | 0.192 | 0.064 | 0.16 | 0.416 | 1.3
97 | 20 | 0.12 | 32.3333 | 0.1 | 0.22 | 0.35567 | 1.1 | 0.194 | 0.064667 | 0.16167 | 0.42033 | 1.3
98 | 20 | 0.12 | 32.6667 | 0.1 | 0.22 | 0.35933 | 1.1 | 0.196 | 0.065333 | 0.16333 | 0.42467 | 1.3
99 | 20 | 0.12 | 33 | 0.1 | 0.22 | 0.363 | 1.1 | 0.198 | 0.066 | 0.165 | 0.429 | 1.3
100 | 20 | 0.12 | 33.3333 | 0.1 | 0.22 | 0.36667 | 1.1 | 0.2 | 0.066667 | 0.16667 | 0.43333 | 1.3
101 | 20 | 0.12 | 33.6667 | 0.1 | 0.22 | 0.37033 | 1.1 | 0.202 | 0.067333 | 0.16833 | 0.43767 | 1.3
102 | 20 | 0.12 | 34 | 0.1 | 0.22 | 0.374 | 1.1 | 0.204 | 0.068 | 0.17 | 0.442 | 1.3
103 | 20 | 0.12 | 34.3333 | 0.1 | 0.22 | 0.37767 | 1.1 | 0.206 | 0.068667 | 0.17167 | 0.44633 | 1.3
104 | 20 | 0.12 | 34.6667 | 0.1 | 0.22 | 0.38133 | 1.1 | 0.208 | 0.069333 | 0.17333 | 0.45067 | 1.3
105 | 20 | 0.12 | 35 | 0.1 | 0.22 | 0.385 | 1.1 | 0.21 | 0.07 | 0.175 | 0.455 | 1.3
106 | 20 | 0.12 | 35.3333 | 0.1 | 0.22 | 0.38867 | 1.1 | 0.212 | 0.070667 | 0.17667 | 0.45933 | 1.3
107 | 20 | 0.12 | 35.6667 | 0.1 | 0.22 | 0.39233 | 1.1 | 0.214 | 0.071333 | 0.17833 | 0.46367 | 1.3
108 | 20 | 0.12 | 36 | 0.1 | 0.22 | 0.396 | 1.1 | 0.216 | 0.072 | 0.18 | 0.468 | 1.3
109 | 20 | 0.12 | 36.3333 | 0.1 | 0.22 | 0.39967 | 1.1 | 0.218 | 0.072667 | 0.18167 | 0.47233 | 1.3
110 | 20 | 0.12 | 36.6667 | 0.1 | 0.22 | 0.40333 | 1.1 | 0.22 | 0.073333 | 0.18333 | 0.47667 | 1.3
111 | 20 | 0.12 | 37 | 0.1 | 0.22 | 0.407 | 1.1 | 0.222 | 0.074 | 0.185 | 0.481 | 1.3
112 | 20 | 0.12 | 37.3333 | 0.1 | 0.22 | 0.41067 | 1.1 | 0.224 | 0.074667 | 0.18667 | 0.48533 | 1.3
113 | 20 | 0.12 | 37.6667 | 0.1 | 0.22 | 0.41433 | 1.1 | 0.226 | 0.075333 | 0.18833 | 0.48967 | 1.3
114 | 20 | 0.12 | 38 | 0.1 | 0.22 | 0.418 | 1.1 | 0.228 | 0.076 | 0.19 | 0.494 | 1.3
115 | 20 | 0.12 | 38.3333 | 0.1 | 0.22 | 0.42167 | 1.1 | 0.23 | 0.076667 | 0.19167 | 0.49833 | 1.3
116 | 20 | 0.12 | 38.6667 | 0.1 | 0.22 | 0.42533 | 1.1 | 0.232 | 0.077333 | 0.19333 | 0.50267 | 1.3
117 | 20 | 0.12 | 39 | 0.1 | 0.22 | 0.429 | 1.1 | 0.234 | 0.078 | 0.195 | 0.507 | 1.3
118 | 20 | 0.12 | 39.3333 | 0.1 | 0.22 | 0.43267 | 1.1 | 0.236 | 0.078667 | 0.19667 | 0.51133 | 1.3
119 | 20 | 0.12 | 39.6667 | 0.1 | 0.22 | 0.43633 | 1.1 | 0.238 | 0.079333 | 0.19833 | 0.51567 | 1.3
120 | 20 | 0.12 | 40 | 0.1 | 0.22 | 0.44 | 1.1 | 0.24 | 0.08 | 0.2 | 0.52 | 1.3
121 | 20 | 0.12 | 40.3333 | 0.1 | 0.22 | 0.44367 | 1.1 | 0.242 | 0.080667 | 0.20167 | 0.52433 | 1.3
122 | 20 | 0.12 | 40.6667 | 0.1 | 0.22 | 0.44733 | 1.1 | 0.244 | 0.081333 | 0.20333 | 0.52867 | 1.3
123 | 20 | 0.12 | 41 | 0.1 | 0.22 | 0.451 | 1.1 | 0.246 | 0.082 | 0.205 | 0.533 | 1.3
124 | 20 | 0.12 | 41.3333 | 0.1 | 0.22 | 0.45467 | 1.1 | 0.248 | 0.082667 | 0.20667 | 0.53733 | 1.3
125 | 20 | 0.12 | 41.6667 | 0.1 | 0.22 | 0.45833 | 1.1 | 0.25 | 0.083333 | 0.20833 | 0.54167 | 1.3
126 | 20 | 0.12 | 42 | 0.1 | 0.22 | 0.462 | 1.1 | 0.252 | 0.084 | 0.21 | 0.546 | 1.3
127 | 20 | 0.12 | 42.3333 | 0.1 | 0.22 | 0.46567 | 1.1 | 0.254 | 0.084667 | 0.21167 | 0.55033 | 1.3
128 | 20 | 0.12 | 42.6667 | 0.1 | 0.22 | 0.46933 | 1.1 | 0.256 | 0.085333 | 0.21333 | 0.55467 | 1.3
129 | 20 | 0.12 | 43 | 0.1 | 0.22 | 0.473 | 1.1 | 0.258 | 0.086 | 0.215 | 0.559 | 1.3
130 | 20 | 0.12 | 43.3333 | 0.1 | 0.22 | 0.47667 | 1.1 | 0.26 | 0.086667 | 0.21667 | 0.56333 | 1.3
131 | 20 | 0.12 | 43.6667 | 0.1 | 0.22 | 0.48033 | 1.1 | 0.262 | 0.087333 | 0.21833 | 0.56767 | 1.3
132 | 20 | 0.12 | 44 | 0.1 | 0.22 | 0.484 | 1.1 | 0.264 | 0.088 | 0.22 | 0.572 | 1.3
133 | 20 | 0.12 | 44.3333 | 0.1 | 0.22 | 0.48767 | 1.1 | 0.266 | 0.088667 | 0.22167 | 0.57633 | 1.3
134 | 20 | 0.12 | 44.6667 | 0.1 | 0.22 | 0.49133 | 1.1 | 0.268 | 0.089333 | 0.22333 | 0.58067 | 1.3
135 | 20 | 0.12 | 45 | 0.1 | 0.22 | 0.495 | 1.1 | 0.27 | 0.09 | 0.225 | 0.585 | 1.3
136 | 20 | 0.12 | 45.3333 | 0.1 | 0.22 | 0.49867 | 1.1 | 0.272 | 0.090667 | 0.22667 | 0.58933 | 1.3
137 | 20 | 0.12 | 45.6667 | 0.1 | 0.22 | 0.50233 | 1.1 | 0.274 | 0.091333 | 0.22833 | 0.59367 | 1.3
138 | 20 | 0.12 | 46 | 0.1 | 0.22 | 0.506 | 1.1 | 0.276 | 0.092 | 0.23 | 0.598 | 1.3
139 | 20 | 0.12 | 46.3333 | 0.1 | 0.22 | 0.50967 | 1.1 | 0.278 | 0.092667 | 0.23167 | 0.60233 | 1.3
140 | 20 | 0.12 | 46.6667 | 0.1 | 0.22 | 0.51333 | 1.1 | 0.28 | 0.093333 | 0.23333 | 0.60667 | 1.3
141 | 20 | 0.12 | 47 | 0.1 | 0.22 | 0.517 | 1.1 | 0.282 | 0.094 | 0.235 | 0.611 | 1.3
142 | 20 | 0.12 | 47.3333 | 0.1 | 0.22 | 0.52067 | 1.1 | 0.284 | 0.094667 | 0.23667 | 0.61533 | 1.3
143 | 20 | 0.12 | 47.6667 | 0.1 | 0.22 | 0.52433 | 1.1 | 0.286 | 0.095333 | 0.23833 | 0.61967 | 1.3
144 | 20 | 0.12 | 48 | 0.1 | 0.22 | 0.528 | 1.1 | 0.288 | 0.096 | 0.24 | 0.624 | 1.3
145 | 20 | 0.12 | 48.3333 | 0.1 | 0.22 | 0.53167 | 1.1 | 0.29 | 0.096667 | 0.24167 | 0.62833 | 1.3
146 | 20 | 0.12 | 48.6667 | 0.1 | 0.22 | 0.53533 | 1.1 | 0.292 | 0.097333 | 0.24333 | 0.63267 | 1.3
147 | 20 | 0.12 | 49 | 0.1 | 0.22 | 0.539 | 1.1 | 0.294 | 0.098 | 0.245 | 0.637 | 1.3
148 | 20 | 0.12 | 49.3333 | 0.1 | 0.22 | 0.54267 | 1.1 | 0.296 | 0.098667 | 0.24667 | 0.64133 | 1.3
149 | 20 | 0.12 | 49.6667 | 0.1 | 0.22 | 0.54633 | 1.1 | 0.298 | 0.099333 | 0.24833 | 0.64567 | 1.3
150 | 20 | 0.12 | 50 | 0.1 | 0.22 | 0.55 | 1.1 | 0.3 | 0.1 | 0.25 | 0.65 | 1.3
151 | 20 | 0.12 | 50.3333 | 0.1 | 0.22 | 0.55367 | 1.1 | 0.302 | 0.10067 | 0.25167 | 0.65433 | 1.3
152 | 20 | 0.12 | 50.6667 | 0.1 | 0.22 | 0.55733 | 1.1 | 0.304 | 0.10133 | 0.25333 | 0.65867 | 1.3
153 | 20 | 0.12 | 51 | 0.1 | 0.22 | 0.561 | 1.1 | 0.306 | 0.102 | 0.255 | 0.663 | 1.3
154 | 20 | 0.12 | 51.3333 | 0.1 | 0.22 | 0.56467 | 1.1 | 0.308 | 0.10267 | 0.25667 | 0.66733 | 1.3
155 | 20 | 0.12 | 51.6667 | 0.1 | 0.22 | 0.56833 | 1.1 | 0.31 | 0.10333 | 0.25833 | 0.67167 | 1.3
156 | 20 | 0.12 | 52 | 0.1 | 0.22 | 0.572 | 1.1 | 0.312 | 0.104 | 0.26 | 0.676 | 1.3
157 | 20 | 0.12 | 52.3333 | 0.1 | 0.22 | 0.57567 | 1.1 | 0.314 | 0.10467 | 0.26167 | 0.68033 | 1.3
158 | 20 | 0.12 | 52.6667 | 0.1 | 0.22 | 0.57933 | 1.1 | 0.316 | 0.10533 | 0.26333 | 0.68467 | 1.3
159 | 20 | 0.12 | 53 | 0.1 | 0.22 | 0.583 | 1.1 | 0.318 | 0.106 | 0.265 | 0.689 | 1.3
160 | 20 | 0.12 | 53.3333 | 0.1 | 0.22 | 0.58667 | 1.1 | 0.32 | 0.10667 | 0.26667 | 0.69333 | 1.3
161 | 20 | 0.12 | 53.6667 | 0.1 | 0.22 | 0.59033 | 1.1 | 0.322 | 0.10733 | 0.26833 | 0.69767 | 1.3
162 | 20 | 0.12 | 54 | 0.1 | 0.22 | 0.594 | 1.1 | 0.324 | 0.108 | 0.27 | 0.702 | 1.3
163 | 20 | 0.12 | 54.3333 | 0.1 | 0.22 | 0.59767 | 1.1 | 0.326 | 0.10867 | 0.27167 | 0.70633 | 1.3
164 | 20 | 0.12 | 54.6667 | 0.1 | 0.22 | 0.60133 | 1.1 | 0.328 | 0.10933 | 0.27333 | 0.71067 | 1.3
165 | 20 | 0.12 | 55 | 0.1 | 0.22 | 0.605 | 1.1 | 0.33 | 0.11 | 0.275 | 0.715 | 1.3
166 | 20 | 0.12 | 55.3333 | 0.1 | 0.22 | 0.60867 | 1.1 | 0.332 | 0.11067 | 0.27667 | 0.71933 | 1.3
167 | 20 | 0.12 | 55.6667 | 0.1 | 0.22 | 0.61233 | 1.1 | 0.334 | 0.11133 | 0.27833 | 0.72367 | 1.3
168 | 20 | 0.12 | 56 | 0.1 | 0.22 | 0.616 | 1.1 | 0.336 | 0.112 | 0.28 | 0.728 | 1.3
169 | 20 | 0.12 | 56.3333 | 0.1 | 0.22 | 0.61967 | 1.1 | 0.338 | 0.11267 | 0.28167 | 0.73233 | 1.3
170 | 20 | 0.12 | 56.6667 | 0.1 | 0.22 | 0.62333 | 1.1 | 0.34 | 0.11333 | 0.28333 | 0.73667 | 1.3
171 | 20 | 0.12 | 57 | 0.1 | 0.22 | 0.627 | 1.1 | 0.342 | 0.114 | 0.285 | 0.741 | 1.3
172 | 20 | 0.12 | 57.3333 | 0.1 | 0.22 | 0.63067 | 1.1 | 0.344 | 0.11467 | 0.28667 | 0.74533 | 1.3
173 | 20 | 0.12 | 57.6667 | 0.1 | 0.22 | 0.63433 | 1.1 | 0.346 | 0.11533 | 0.28833 | 0.74967 | 1.3
174 | 20 | 0.12 | 58 | 0.1 | 0.22 | 0.638 | 1.1 | 0.348 | 0.116 | 0.29 | 0.754 | 1.3
175 | 20 | 0.12 | 58.3333 | 0.1 | 0.22 | 0.64167 | 1.1 | 0.35 | 0.11667 | 0.29167 | 0.75833 | 1.3
176 | 20 | 0.12 | 58.6667 | 0.1 | 0.22 | 0.64533 | 1.1 | 0.352 | 0.11733 | 0.29333 | 0.76267 | 1.3
177 | 20 | 0.12 | 59 | 0.1 | 0.22 | 0.649 | 1.1 | 0.354 | 0.118 | 0.295 | 0.767 | 1.3
178 | 20 | 0.12 | 59.3333 | 0.1 | 0.22 | 0.65267 | 1.1 | 0.356 | 0.11867 | 0.29667 | 0.77133 | 1.3
179 | 20 | 0.12 | 59.6667 | 0.1 | 0.22 | 0.65633 | 1.1 | 0.358 | 0.11933 | 0.29833 | 0.77567 | 1.3
180 | 20 | 0.12 | 60 | 0.1 | 0.22 | 0.66 | 1.1 | 0.36 | 0.12 | 0.3 | 0.78 | 1.3
181 | 20 | 0.12 | 60.3333 | 0.1 | 0.22 | 0.66367 | 1.1 | 0.362 | 0.12067 | 0.30167 | 0.78433 | 1.3
182 | 20 | 0.12 | 60.6667 | 0.1 | 0.22 | 0.66733 | 1.1 | 0.364 | 0.12133 | 0.30333 | 0.78867 | 1.3
183 | 20 | 0.12 | 61 | 0.1 | 0.22 | 0.671 | 1.1 | 0.366 | 0.122 | 0.305 | 0.793 | 1.3
184 | 20 | 0.12 | 61.3333 | 0.1 | 0.22 | 0.67467 | 1.1 | 0.368 | 0.12267 | 0.30667 | 0.79733 | 1.3
185 | 20 | 0.12 | 61.6667 | 0.1 | 0.22 | 0.67833 | 1.1 | 0.37 | 0.12333 | 0.30833 | 0.80167 | 1.3
186 | 20 | 0.12 | 62 | 0.1 | 0.22 | 0.682 | 1.1 | 0.372 | 0.124 | 0.31 | 0.806 | 1.3
187 | 20 | 0.12 | 62.3333 | 0.1 | 0.22 | 0.68567 | 1.1 | 0.374 | 0.12467 | 0.31167 | 0.81033 | 1.3
188 | 20 | 0.12 | 62.6667 | 0.1 | 0.22 | 0.68933 | 1.1 | 0.376 | 0.12533 | 0.31333 | 0.81467 | 1.3
189 | 20 | 0.12 | 63 | 0.1 | 0.22 | 0.693 | 1.1 | 0.378 | 0.126 | 0.315 | 0.819 | 1.3
190 | 20 | 0.12 | 63.3333 | 0.1 | 0.22 | 0.69667 | 1.1 | 0.38 | 0.12667 | 0.31667 | 0.82333 | 1.3
191 | 20 | 0.12 | 63.6667 | 0.1 | 0.22 | 0.70033 | 1.1 | 0.382 | 0.12733 | 0.31833 | 0.82767 | 1.3
192 | 20 | 0.12 | 64 | 0.1 | 0.22 | 0.704 | 1.1 | 0.384 | 0.128 | 0.32 | 0.832 | 1.3
193 | 20 | 0.12 | 64.3333 | 0.1 | 0.22 | 0.70767 | 1.1 | 0.386 | 0.12867 | 0.32167 | 0.83633 | 1.3
194 | 20 | 0.12 | 64.6667 | 0.1 | 0.22 | 0.71133 | 1.1 | 0.388 | 0.12933 | 0.32333 | 0.84067 | 1.3
195 | 20 | 0.12 | 65 | 0.1 | 0.22 | 0.715 | 1.1 | 0.39 | 0.13 | 0.325 | 0.845 | 1.3
196 | 20 | 0.12 | 65.3333 | 0.1 | 0.22 | 0.71867 | 1.1 | 0.392 | 0.13067 | 0.32667 | 0.84933 | 1.3
197 | 20 | 0.12 | 65.6667 | 0.1 | 0.22 | 0.72233 | 1.1 | 0.394 | 0.13133 | 0.32833 | 0.85367 | 1.3
198 | 20 | 0.12 | 66 | 0.1 | 0.22 | 0.726 | 1.1 | 0.396 | 0.132 | 0.33 | 0.858 | 1.3
199 | 20 | 0.12 | 66.3333 | 0.1 | 0.22 | 0.72967 | 1.1 | 0.398 | 0.13267 | 0.33167 | 0.86233 | 1.3
200 | 20 | 0.12 | 66.6667 | 0.1 | 0.22 | 0.73333 | 1.1 | 0.4 | 0.13333 | 0.33333 | 0.86667 | 1.3
201 | 20 | 0.12 | 67 | 0.1 | 0.22 | 0.737 | 1.1 | 0.402 | 0.134 | 0.335 | 0.871 | 1.3
202 | 20 | 0.12 | 67.3333 | 0.1 | 0.22 | 0.74067 | 1.1 | 0.404 | 0.13467 | 0.33667 | 0.87533 | 1.3
203 | 20 | 0.12 | 67.6667 | 0.1 | 0.22 | 0.74433 | 1.1 | 0.406 | 0.13533 | 0.33833 | 0.87967 | 1.3
204 | 20 | 0.12 | 68 | 0.1 | 0.22 | 0.748 | 1.1 | 0.408 | 0.136 | 0.34 | 0.884 | 1.3
205 | 20 | 0.12 | 68.3333 | 0.1 | 0.22 | 0.75167 | 1.1 | 0.41 | 0.13667 | 0.34167 | 0.88833 | 1.3
206 | 20 | 0.12 | 68.6667 | 0.1 | 0.22 | 0.75533 | 1.1 | 0.412 | 0.13733 | 0.34333 | 0.89267 | 1.3
207 | 20 | 0.12 | 69 | 0.1 | 0.22 | 0.759 | 1.1 | 0.414 | 0.138 | 0.345 | 0.897 | 1.3
208 | 20 | 0.12 | 69.3333 | 0.1 | 0.22 | 0.76267 | 1.1 | 0.416 | 0.13867 | 0.34667 | 0.90133 | 1.3
209 | 20 | 0.12 | 69.6667 | 0.1 | 0.22 | 0.76633 | 1.1 | 0.418 | 0.13933 | 0.34833 | 0.90567 | 1.3
210 | 20 | 0.12 | 70 | 0.1 | 0.22 | 0.77 | 1.1 | 0.42 | 0.14 | 0.35 | 0.91 | 1.3
211 | 20 | 0.12 | 70.3333 | 0.1 | 0.22 | 0.77367 | 1.1 | 0.422 | 0.14067 | 0.35167 | 0.91433 | 1.3
212 | 20 | 0.12 | 70.6667 | 0.1 | 0.22 | 0.77733 | 1.1 | 0.424 | 0.14133 | 0.35333 | 0.91867 | 1.3
213 | 20 | 0.12 | 71 | 0.1 | 0.22 | 0.781 | 1.1 | 0.426 | 0.142 | 0.355 | 0.923 | 1.3
214 | 20 | 0.12 | 71.3333 | 0.1 | 0.22 | 0.78467 | 1.1 | 0.428 | 0.14267 | 0.35667 | 0.92733 | 1.3
215 | 20 | 0.12 | 71.6667 | 0.1 | 0.22 | 0.78833 | 1.1 | 0.43 | 0.14333 | 0.35833 | 0.93167 | 1.3
216 | 20 | 0.12 | 72 | 0.1 | 0.22 | 0.792 | 1.1 | 0.432 | 0.144 | 0.36 | 0.936 | 1.3
217 | 20 | 0.12 | 72.3333 | 0.1 | 0.22 | 0.79567 | 1.1 | 0.434 | 0.14467 | 0.36167 | 0.94033 | 1.3
218 | 20 | 0.12 | 72.6667 | 0.1 | 0.22 | 0.79933 | 1.1 | 0.436 | 0.14533 | 0.36333 | 0.94467 | 1.3
219 | 20 | 0.12 | 73 | 0.1 | 0.22 | 0.803 | 1.1 | 0.438 | 0.146 | 0.365 | 0.949 | 1.3
220 | 20 | 0.12 | 73.3333 | 0.1 | 0.22 | 0.80667 | 1.1 | 0.44 | 0.14667 | 0.36667 | 0.95333 | 1.3
221 | 20 | 0.12 | 73.6667 | 0.1 | 0.22 | 0.81033 | 1.1 | 0.442 | 0.14733 | 0.36833 | 0.95767 | 1.3
222 | 20 | 0.12 | 74 | 0.1 | 0.22 | 0.814 | 1.1 | 0.444 | 0.148 | 0.37 | 0.962 | 1.3
223 | 20 | 0.12 | 74.3333 | 0.1 | 0.22 | 0.81767 | 1.1 | 0.446 | 0.14867 | 0.37167 | 0.96633 | 1.3
224 | 20 | 0.12 | 74.6667 | 0.1 | 0.22 | 0.82133 | 1.1 | 0.448 | 0.14933 | 0.37333 | 0.97067 | 1.3
225 | 20 | 0.12 | 75 | 0.1 | 0.22 | 0.825 | 1.1 | 0.45 | 0.15 | 0.375 | 0.975 | 1.3
226 | 20 | 0.12 | 75.3333 | 0.1 | 0.22 | 0.82867 | 1.1 | 0.452 | 0.15067 | 0.37667 | 0.97933 | 1.3
227 | 20 | 0.12 | 75.6667 | 0.1 | 0.22 | 0.83233 | 1.1 | 0.454 | 0.15133 | 0.37833 | 0.98367 | 1.3
228 | 20 | 0.12 | 76 | 0.1 | 0.22 | 0.836 | 1.1 | 0.456 | 0.152 | 0.38 | 0.988 | 1.3
229 | 20 | 0.12 | 76.3333 | 0.1 | 0.22 | 0.83967 | 1.1 | 0.458 | 0.15267 | 0.38167 | 0.99233 | 1.3
230 | 20 | 0.12 | 76.6667 | 0.1 | 0.22 | 0.84333 | 1.1 | 0.46 | 0.15333 | 0.38333 | 0.99667 | 1.3
231 | 20 | 0.12 | 77 | 0.1 | 0.22 | 0.847 | 1.1 | 0.462 | 0.154 | 0.385 | 1.001 | 1.3
232 | 20 | 0.12 | 77.3333 | 0.1 | 0.22 | 0.85067 | 1.1 | 0.464 | 0.15467 | 0.38667 | 1.0053 | 1.3
233 | 20 | 0.12 | 77.6667 | 0.1 | 0.22 | 0.85433 | 1.1 | 0.466 | 0.15533 | 0.38833 | 1.0097 | 1.3
234 | 20 | 0.12 | 78 | 0.1 | 0.22 | 0.858 | 1.1 | 0.468 | 0.156 | 0.39 | 1.014 | 1.3
235 | 20 | 0.12 | 78.3333 | 0.1 | 0.22 | 0.86167 | 1.1 | 0.47 | 0.15667 | 0.39167 | 1.0183 | 1.3
236 | 20 | 0.12 | 78.6667 | 0.1 | 0.22 | 0.86533 | 1.1 | 0.472 | 0.15733 | 0.39333 | 1.0227 | 1.3
237 | 20 | 0.12 | 79 | 0.1 | 0.22 | 0.869 | 1.1 | 0.474 | 0.158 | 0.395 | 1.027 | 1.3
238 | 20 | 0.12 | 79.3333 | 0.1 | 0.22 | 0.87267 | 1.1 | 0.476 | 0.15867 | 0.39667 | 1.0313 | 1.3
239 | 20 | 0.12 | 79.6667 | 0.1 | 0.22 | 0.87633 | 1.1 | 0.478 | 0.15933 | 0.39833 | 1.0357 | 1.3
240 | 20 | 0.12 | 80 | 0.1 | 0.22 | 0.88 | 1.1 | 0.48 | 0.16 | 0.4 | 1.04 | 1.3
241 | 20 | 0.12 | 80.3333 | 0.1 | 0.22 | 0.88367 | 1.1 | 0.482 | 0.16067 | 0.40167 | 1.0443 | 1.3
242 | 20 | 0.12 | 80.6667 | 0.1 | 0.22 | 0.88733 | 1.1 | 0.484 | 0.16133 | 0.40333 | 1.0487 | 1.3
243 | 20 | 0.12 | 81 | 0.1 | 0.22 | 0.891 | 1.1 | 0.486 | 0.162 | 0.405 | 1.053 | 1.3
244 | 20 | 0.12 | 81.3333 | 0.1 | 0.22 | 0.89467 | 1.1 | 0.488 | 0.16267 | 0.40667 | 1.0573 | 1.3
245 | 20 | 0.12 | 81.6667 | 0.1 | 0.22 | 0.89833 | 1.1 | 0.49 | 0.16333 | 0.40833 | 1.0617 | 1.3
246 | 20 | 0.12 | 82 | 0.1 | 0.22 | 0.902 | 1.1 | 0.492 | 0.164 | 0.41 | 1.066 | 1.3
247 | 20 | 0.12 | 82.3333 | 0.1 | 0.22 | 0.90567 | 1.1 | 0.494 | 0.16467 | 0.41167 | 1.0703 | 1.3
248 | 20 | 0.12 | 82.6667 | 0.1 | 0.22 | 0.90933 | 1.1 | 0.496 | 0.16533 | 0.41333 | 1.0747 | 1.3
249 | 20 | 0.12 | 83 | 0.1 | 0.22 | 0.913 | 1.1 | 0.498 | 0.166 | 0.415 | 1.079 | 1.3
250 | 20 | 0.12 | 83.3333 | 0.1 | 0.22 | 0.91667 | 1.1 | 0.5 | 0.16667 | 0.41667 | 1.0833 | 1.3
251 | 20 | 0.12 | 83.6667 | 0.1 | 0.22 | 0.92033 | 1.1 | 0.502 | 0.16733 | 0.41833 | 1.0877 | 1.3
252 | 20 | 0.12 | 84 | 0.1 | 0.22 | 0.924 | 1.1 | 0.504 | 0.168 | 0.42 | 1.092 | 1.3
253 | 20 | 0.12 | 84.3333 | 0.1 | 0.22 | 0.92767 | 1.1 | 0.506 | 0.16867 | 0.42167 | 1.0963 | 1.3
254 | 20 | 0.12 | 84.6667 | 0.1 | 0.22 | 0.93133 | 1.1 | 0.508 | 0.16933 | 0.42333 | 1.1007 | 1.3
255 | 20 | 0.12 | 85 | 0.1 | 0.22 | 0.935 | 1.1 | 0.51 | 0.17 | 0.425 | 1.105 | 1.3
256 | 20 | 0.12 | 85.3333 | 0.1 | 0.22 | 0.93867 | 1.1 | 0.512 | 0.17067 | 0.42667 | 1.1093 | 1.3
257 | 20 | 0.12 | 85.6667 | 0.1 | 0.22 | 0.94233 | 1.1 | 0.514 | 0.17133 | 0.42833 | 1.1137 | 1.3
258 | 20 | 0.12 | 86 | 0.1 | 0.22 | 0.946 | 1.1 | 0.516 | 0.172 | 0.43 | 1.118 | 1.3
259 | 20 | 0.12 | 86.3333 | 0.1 | 0.22 | 0.94967 | 1.1 | 0.518 | 0.17267 | 0.43167 | 1.1223 | 1.3
260 | 20 | 0.12 | 86.6667 | 0.1 | 0.22 | 0.95333 | 1.1 | 0.52 | 0.17333 | 0.43333 | 1.1267 | 1.3
261 | 20 | 0.12 | 87 | 0.1 | 0.22 | 0.957 | 1.1 | 0.522 | 0.174 | 0.435 | 1.131 | 1.3
262 | 20 | 0.12 | 87.3333 | 0.1 | 0.22 | 0.96067 | 1.1 | 0.524 | 0.17467 | 0.43667 | 1.1353 | 1.3
263 | 20 | 0.12 | 87.6667 | 0.1 | 0.22 | 0.96433 | 1.1 | 0.526 | 0.17533 | 0.43833 | 1.1397 | 1.3
264 | 20 | 0.12 | 88 | 0.1 | 0.22 | 0.968 | 1.1 | 0.528 | 0.176 | 0.44 | 1.144 | 1.3
265 | 20 | 0.12 | 88.3333 | 0.1 | 0.22 | 0.97167 | 1.1 | 0.53 | 0.17667 | 0.44167 | 1.1483 | 1.3
266 | 20 | 0.12 | 88.6667 | 0.1 | 0.22 | 0.97533 | 1.1 | 0.532 | 0.17733 | 0.44333 | 1.1527 | 1.3
267 | 20 | 0.12 | 89 | 0.1 | 0.22 | 0.979 | 1.1 | 0.534 | 0.178 | 0.445 | 1.157 | 1.3
268 | 20 | 0.12 | 89.3333 | 0.1 | 0.22 | 0.98267 | 1.1 | 0.536 | 0.17867 | 0.44667 | 1.1613 | 1.3
269 | 20 | 0.12 | 89.6667 | 0.1 | 0.22 | 0.98633 | 1.1 | 0.538 | 0.17933 | 0.44833 | 1.1657 | 1.3
270 | 20 | 0.12 | 90 | 0.1 | 0.22 | 0.99 | 1.1 | 0.54 | 0.18 | 0.45 | 1.17 | 1.3
271 | 20 | 0.12 | 90.3333 | 0.1 | 0.22 | 0.99367 | 1.1 | 0.542 | 0.18067 | 0.45167 | 1.1743 | 1.3
272 | 20 | 0.12 | 90.6667 | 0.1 | 0.22 | 0.99733 | 1.1 | 0.544 | 0.18133 | 0.45333 | 1.1787 | 1.3
273 | 20 | 0.12 | 91 | 0.1 | 0.22 | 1.001 | 1.1 | 0.546 | 0.182 | 0.455 | 1.183 | 1.3
274 | 20 | 0.12 | 91.3333 | 0.1 | 0.22 | 1.0047 | 1.1 | 0.548 | 0.18267 | 0.45667 | 1.1873 | 1.3
275 | 20 | 0.12 | 91.6667 | 0.1 | 0.22 | 1.0083 | 1.1 | 0.55 | 0.18333 | 0.45833 | 1.1917 | 1.3
276 | 20 | 0.12 | 92 | 0.1 | 0.22 | 1.012 | 1.1 | 0.552 | 0.184 | 0.46 | 1.196 | 1.3
277 | 20 | 0.12 | 92.3333 | 0.1 | 0.22 | 1.0157 | 1.1 | 0.554 | 0.18467 | 0.46167 | 1.2003 | 1.3
278 | 20 | 0.12 | 92.6667 | 0.1 | 0.22 | 1.0193 | 1.1 | 0.556 | 0.18533 | 0.46333 | 1.2047 | 1.3
279 | 20 | 0.12 | 93 | 0.1 | 0.22 | 1.023 | 1.1 | 0.558 | 0.186 | 0.465 | 1.209 | 1.3
280 | 20 | 0.12 | 93.3333 | 0.1 | 0.22 | 1.0267 | 1.1 | 0.56 | 0.18667 | 0.46667 | 1.2133 | 1.3
281 | 20 | 0.12 | 93.6667 | 0.1 | 0.22 | 1.0303 | 1.1 | 0.562 | 0.18733 | 0.46833 | 1.2177 | 1.3
282 | 20 | 0.12 | 94 | 0.1 | 0.22 | 1.034 | 1.1 | 0.564 | 0.188 | 0.47 | 1.222 | 1.3
283 | 20 | 0.12 | 94.3333 | 0.1 | 0.22 | 1.0377 | 1.1 | 0.566 | 0.18867 | 0.47167 | 1.2263 | 1.3
284 | 20 | 0.12 | 94.6667 | 0.1 | 0.22 | 1.0413 | 1.1 | 0.568 | 0.18933 | 0.47333 | 1.2307 | 1.3
285 | 20 | 0.12 | 95 | 0.1 | 0.22 | 1.045 | 1.1 | 0.57 | 0.19 | 0.475 | 1.235 | 1.3
286 | 20 | 0.12 | 95.3333 | 0.1 | 0.22 | 1.0487 | 1.1 | 0.572 | 0.19067 | 0.47667 | 1.2393 | 1.3
287 | 20 | 0.12 | 95.6667 | 0.1 | 0.22 | 1.0523 | 1.1 | 0.574 | 0.19133 | 0.47833 | 1.2437 | 1.3
288 | 20 | 0.12 | 96 | 0.1 | 0.22 | 1.056 | 1.1 | 0.576 | 0.192 | 0.48 | 1.248 | 1.3
289 | 20 | 0.12 | 96.3333 | 0.1 | 0.22 | 1.0597 | 1.1 | 0.578 | 0.19267 | 0.48167 | 1.2523 | 1.3
290 | 20 | 0.12 | 96.6667 | 0.1 | 0.22 | 1.0633 | 1.1 | 0.58 | 0.19333 | 0.48333 | 1.2567 | 1.3
291 | 20 | 0.12 | 97 | 0.1 | 0.22 | 1.067 | 1.1 | 0.582 | 0.194 | 0.485 | 1.261 | 1.3
292 | 20 | 0.12 | 97.3333 | 0.1 | 0.22 | 1.0707 | 1.1 | 0.584 | 0.19467 | 0.48667 | 1.2653 | 1.3
293 | 20 | 0.12 | 97.6667 | 0.1 | 0.22 | 1.0743 | 1.1 | 0.586 | 0.19533 | 0.48833 | 1.2697 | 1.3
294 | 20 | 0.12 | 98 | 0.1 | 0.22 | 1.078 | 1.1 | 0.588 | 0.196 | 0.49 | 1.274 | 1.3
295 | 20 | 0.12 | 98.3333 | 0.1 | 0.22 | 1.0817 | 1.1 | 0.59 | 0.19667 | 0.49167 | 1.2783 | 1.3
296 | 20 | 0.12 | 98.6667 | 0.1 | 0.22 | 1.0853 | 1.1 | 0.592 | 0.19733 | 0.49333 | 1.2827 | 1.3
297 | 20 | 0.12 | 99 | 0.1 | 0.22 | 1.089 | 1.1 | 0.594 | 0.198 | 0.495 | 1.287 | 1.3
298 | 20 | 0.12 | 99.3333 | 0.1 | 0.22 | 1.0927 | 1.1 | 0.596 | 0.19867 | 0.49667 | 1.2913 | 1.3
299 | 20 | 0.12 | 99.6667 | 0.1 | 0.22 | 1.0963 | 1.1 | 0.598 | 0.19933 | 0.49833 | 1.2957 | 1.3
300 | 20 | 0.12 | 100 | 0.1 | 0.22 | 1.1 | 1.1 | 0.6 | 0.2 | 0.5 | 1.3 | 1.3
301 | 20 | 0.12 | 100.3333 | 0.1 | 0.22 | 1.1037 | 1.1 | 0.602 | 0.20067 | 0.50167 | 1.3043 | 1.3
302 | 20 | 0.12 | 100.6667 | 0.1 | 0.22 | 1.1073 | 1.1 | 0.604 | 0.20133 | 0.50333 | 1.3087 | 1.3
303 | 20 | 0.12 | 101 | 0.1 | 0.22 | 1.111 | 1.1 | 0.606 | 0.202 | 0.505 | 1.313 | 1.3
304 | 20 | 0.12 | 101.3333 | 0.1 | 0.22 | 1.1147 | 1.1 | 0.608 | 0.20267 | 0.50667 | 1.3173 | 1.3
305 | 20 | 0.12 | 101.6667 | 0.1 | 0.22 | 1.1183 | 1.1 | 0.61 | 0.20333 | 0.50833 | 1.3217 | 1.3
306 | 20 | 0.12 | 102 | 0.1 | 0.22 | 1.122 | 1.1 | 0.612 | 0.204 | 0.51 | 1.326 | 1.3
307 | 20 | 0.12 | 102.3333 | 0.1 | 0.22 | 1.1257 | 1.1 | 0.614 | 0.20467 | 0.51167 | 1.3303 | 1.3
308 | 20 | 0.12 | 102.6667 | 0.1 | 0.22 | 1.1293 | 1.1 | 0.616 | 0.20533 | 0.51333 | 1.3347 | 1.3
309 | 20 | 0.12 | 103 | 0.1 | 0.22 | 1.133 | 1.1 | 0.618 | 0.206 | 0.515 | 1.339 | 1.3
310 | 20 | 0.12 | 103.3333 | 0.1 | 0.22 | 1.1367 | 1.1 | 0.62 | 0.20667 | 0.51667 | 1.3433 | 1.3
311 | 20 | 0.12 | 103.6667 | 0.1 | 0.22 | 1.1403 | 1.1 | 0.622 | 0.20733 | 0.51833 | 1.3477 | 1.3
312 | 20 | 0.12 | 104 | 0.1 | 0.22 | 1.144 | 1.1 | 0.624 | 0.208 | 0.52 | 1.352 | 1.3
313 | 20 | 0.12 | 104.3333 | 0.1 | 0.22 | 1.1477 | 1.1 | 0.626 | 0.20867 | 0.52167 | 1.3563 | 1.3
314 | 20 | 0.12 | 104.6667 | 0.1 | 0.22 | 1.1513 | 1.1 | 0.628 | 0.20933 | 0.52333 | 1.3607 | 1.3
315 | 20 | 0.12 | 105 | 0.1 | 0.22 | 1.155 | 1.1 | 0.63 | 0.21 | 0.525 | 1.365 | 1.3
316 | 20 | 0.12 | 105.3333 | 0.1 | 0.22 | 1.1587 | 1.1 | 0.632 | 0.21067 | 0.52667 | 1.3693 | 1.3
317 | 20 | 0.12 | 105.6667 | 0.1 | 0.22 | 1.1623 | 1.1 | 0.634 | 0.21133 | 0.52833 | 1.3737 | 1.3
318 | 20 | 0.12 | 106 | 0.1 | 0.22 | 1.166 | 1.1 | 0.636 | 0.212 | 0.53 | 1.378 | 1.3
319 | 20 | 0.12 | 106.3333 | 0.1 | 0.22 | 1.1697 | 1.1 | 0.638 | 0.21267 | 0.53167 | 1.3823 | 1.3
320 | 20 | 0.12 | 106.6667 | 0.1 | 0.22 | 1.1733 | 1.1 | 0.64 | 0.21333 | 0.53333 | 1.3867 | 1.3
321 | 20 | 0.12 | 107 | 0.1 | 0.22 | 1.177 | 1.1 | 0.642 | 0.214 | 0.535 | 1.391 | 1.3
322 | 20 | 0.12 | 107.3333 | 0.1 | 0.22 | 1.1807 | 1.1 | 0.644 | 0.21467 | 0.53667 | 1.3953 | 1.3
323 | 20 | 0.12 | 107.6667 | 0.1 | 0.22 | 1.1843 | 1.1 | 0.646 | 0.21533 | 0.53833 | 1.3997 | 1.3
324 | 20 | 0.12 | 108 | 0.1 | 0.22 | 1.188 | 1.1 | 0.648 | 0.216 | 0.54 | 1.404 | 1.3
325 | 20 | 0.12 | 108.3333 | 0.1 | 0.22 | 1.1917 | 1.1 | 0.65 | 0.21667 | 0.54167 | 1.4083 | 1.3
326 | 20 | 0.12 | 108.6667 | 0.1 | 0.22 | 1.1953 | 1.1 | 0.652 | 0.21733 | 0.54333 | 1.4127 | 1.3
327 | 20 | 0.12 | 109 | 0.1 | 0.22 | 1.199 | 1.1 | 0.654 | 0.218 | 0.545 | 1.417 | 1.3
328 | 20 | 0.12 | 109.3333 | 0.1 | 0.22 | 1.2027 | 1.1 | 0.656 | 0.21867 | 0.54667 | 1.4213 | 1.3
329 | 20 | 0.12 | 109.6667 | 0.1 | 0.22 | 1.2063 | 1.1 | 0.658 | 0.21933 | 0.54833 | 1.4257 | 1.3
330 | 20 | 0.12 | 110 | 0.1 | 0.22 | 1.21 | 1.1 | 0.66 | 0.22 | 0.55 | 1.43 | 1.3
331 | 20 | 0.12 | 110.3333 | 0.1 | 0.22 | 1.2137 | 1.1 | 0.662 | 0.22067 | 0.55167 | 1.4343 | 1.3
332 | 20 | 0.12 | 110.6667 | 0.1 | 0.22 | 1.2173 | 1.1 | 0.664 | 0.22133 | 0.55333 | 1.4387 | 1.3
333 | 20 | 0.12 | 111 | 0.1 | 0.22 | 1.221 | 1.1 | 0.666 | 0.222 | 0.555 | 1.443 | 1.3
334 | 20 | 0.12 | 111.3333 | 0.1 | 0.22 | 1.2247 | 1.1 | 0.668 | 0.22267 | 0.55667 | 1.4473 | 1.3
335 | 20 | 0.12 | 111.6667 | 0.1 | 0.22 | 1.2283 | 1.1 | 0.67 | 0.22333 | 0.55833 | 1.4517 | 1.3
336 | 20 | 0.12 | 112 | 0.1 | 0.22 | 1.232 | 1.1 | 0.672 | 0.224 | 0.56 | 1.456 | 1.3
337 | 20 | 0.12 | 112.3333 | 0.1 | 0.22 | 1.2357 | 1.1 | 0.674 | 0.22467 | 0.56167 | 1.4603 | 1.3
338 | 20 | 0.12 | 112.6667 | 0.1 | 0.22 | 1.2393 | 1.1 | 0.676 | 0.22533 | 0.56333 | 1.4647 | 1.3
339 | 20 | 0.12 | 113 | 0.1 | 0.22 | 1.243 | 1.1 | 0.678 | 0.226 | 0.565 | 1.469 | 1.3
340 | 20 | 0.12 | 113.3333 | 0.1 | 0.22 | 1.2467 | 1.1 | 0.68 | 0.22667 | 0.56667 | 1.4733 | 1.3
341 | 20 | 0.12 | 113.6667 | 0.1 | 0.22 | 1.2503 | 1.1 | 0.682 | 0.22733 | 0.56833 | 1.4777 | 1.3
342 | 20 | 0.12 | 114 | 0.1 | 0.22 | 1.254 | 1.1 | 0.684 | 0.228 | 0.57 | 1.482 | 1.3
343 | 20 | 0.12 | 114.3333 | 0.1 | 0.22 | 1.2577 | 1.1 | 0.686 | 0.22867 | 0.57167 | 1.4863 | 1.3
344 | 20 | 0.12 | 114.6667 | 0.1 | 0.22 | 1.2613 | 1.1 | 0.688 | 0.22933 | 0.57333 | 1.4907 | 1.3
345 | 20 | 0.12 | 115 | 0.1 | 0.22 | 1.265 | 1.1 | 0.69 | 0.23 | 0.575 | 1.495 | 1.3
346 | 20 | 0.12 | 115.3333 | 0.1 | 0.22 | 1.2687 | 1.1 | 0.692 | 0.23067 | 0.57667 | 1.4993 | 1.3
347 | 20 | 0.12 | 115.6667 | 0.1 | 0.22 | 1.2723 | 1.1 | 0.694 | 0.23133 | 0.57833 | 1.5037 | 1.3
348 | 20 | 0.12 | 116 | 0.1 | 0.22 | 1.276 | 1.1 | 0.696 | 0.232 | 0.58 | 1.508 | 1.3
349 | 20 | 0.12 | 116.3333 | 0.1 | 0.22 | 1.2797 | 1.1 | 0.698 | 0.23267 | 0.58167 | 1.5123 | 1.3
350 | 20 | 0.12 | 116.6667 | 0.1 | 0.22 | 1.2833 | 1.1 | 0.7 | 0.23333 | 0.58333 | 1.5167 | 1.3
351 | 20 | 0.12 | 117 | 0.1 | 0.22 | 1.287 | 1.1 | 0.702 | 0.234 | 0.585 | 1.521 | 1.3
352 | 20 | 0.12 | 117.3333 | 0.1 | 0.22 | 1.2907 | 1.1 | 0.704 | 0.23467 | 0.58667 | 1.5253 | 1.3
353 | 20 | 0.12 | 117.6667 | 0.1 | 0.22 | 1.2943 | 1.1 | 0.706 | 0.23533 | 0.58833 | 1.5297 | 1.3
354 | 20 | 0.12 | 118 | 0.1 | 0.22 | 1.298 | 1.1 | 0.708 | 0.236 | 0.59 | 1.534 | 1.3
355 | 20 | 0.12 | 118.3333 | 0.1 | 0.22 | 1.3017 | 1.1 | 0.71 | 0.23667 | 0.59167 | 1.5383 | 1.3
356 | 20 | 0.12 | 118.6667 | 0.1 | 0.22 | 1.3053 | 1.1 | 0.712 | 0.23733 | 0.59333 | 1.5427 | 1.3
357 | 20 | 0.12 | 119 | 0.1 | 0.22 | 1.309 | 1.1 | 0.714 | 0.238 | 0.595 | 1.547 | 1.3
358 | 20 | 0.12 | 119.3333 | 0.1 | 0.22 | 1.3127 | 1.1 | 0.716 | 0.23867 | 0.59667 | 1.5513 | 1.3
359 | 20 | 0.12 | 119.6667 | 0.1 | 0.22 | 1.3163 | 1.1 | 0.718 | 0.23933 | 0.59833 | 1.5557 | 1.3
360 | 20 | 0.12 | 120 | 0.1 | 0.22 | 1.32 | 1.1 | 0.72 | 0.24 | 0.6 | 1.56 | 1.3
361 | 20 | 0.12 | 120.3333 | 0.1 | 0.22 | 1.3237 | 1.1 | 0.722 | 0.24067 | 0.60167 | 1.5643 | 1.3
362 | 20 | 0.12 | 120.6667 | 0.1 | 0.22 | 1.3273 | 1.1 | 0.724 | 0.24133 | 0.60333 | 1.5687 | 1.3
363 | 20 | 0.12 | 121 | 0.1 | 0.22 | 1.331 | 1.1 | 0.726 | 0.242 | 0.605 | 1.573 | 1.3
364 | 20 | 0.12 | 121.3333 | 0.1 | 0.22 | 1.3347 | 1.1 | 0.728 | 0.24267 | 0.60667 | 1.5773 | 1.3
365 | 20 | 0.12 | 121.6667 | 0.1 | 0.22 | 1.3383 | 1.1 | 0.73 | 0.24333 | 0.60833 | 1.5817 | 1.3
366 | 20 | 0.12 | 122 | 0.1 | 0.22 | 1.342 | 1.1 | 0.732 | 0.244 | 0.61 | 1.586 | 1.3
367 | 20 | 0.12 | 122.3333 | 0.1 | 0.22 | 1.3457 | 1.1 | 0.734 | 0.24467 | 0.61167 | 1.5903 | 1.3
368 | 20 | 0.12 | 122.6667 | 0.1 | 0.22 | 1.3493 | 1.1 | 0.736 | 0.24533 | 0.61333 | 1.5947 | 1.3
369 | 20 | 0.12 | 123 | 0.1 | 0.22 | 1.353 | 1.1 | 0.738 | 0.246 | 0.615 | 1.599 | 1.3
370 | 20 | 0.12 | 123.3333 | 0.1 | 0.22 | 1.3567 | 1.1 | 0.74 | 0.24667 | 0.61667 | 1.6033 | 1.3
371 | 20 | 0.12 | 123.6667 | 0.1 | 0.22 | 1.3603 | 1.1 | 0.742 | 0.24733 | 0.61833 | 1.6077 | 1.3
372 | 20 | 0.12 | 124 | 0.1 | 0.22 | 1.364 | 1.1 | 0.744 | 0.248 | 0.62 | 1.612 | 1.3
373 | 20 | 0.12 | 124.3333 | 0.1 | 0.22 | 1.3677 | 1.1 | 0.746 | 0.24867 | 0.62167 | 1.6163 | 1.3
374 | 20 | 0.12 | 124.6667 | 0.1 | 0.22 | 1.3713 | 1.1 | 0.748 | 0.24933 | 0.62333 | 1.6207 | 1.3
375 | 20 | 0.12 | 125 | 0.1 | 0.22 | 1.375 | 1.1 | 0.75 | 0.25 | 0.625 | 1.625 | 1.3
376 | 20 | 0.12 | 125.3333 | 0.1 | 0.22 | 1.3787 | 1.1 | 0.752 | 0.25067 | 0.62667 | 1.6293 | 1.3
377 | 20 | 0.12 | 125.6667 | 0.1 | 0.22 | 1.3823 | 1.1 | 0.754 | 0.25133 | 0.62833 | 1.6337 | 1.3
378 | 20 | 0.12 | 126 | 0.1 | 0.22 | 1.386 | 1.1 | 0.756 | 0.252 | 0.63 | 1.638 | 1.3
379 | 20 | 0.12 | 126.3333 | 0.1 | 0.22 | 1.3897 | 1.1 | 0.758 | 0.25267 | 0.63167 | 1.6423 | 1.3
380 | 20 | 0.12 | 126.6667 | 0.1 | 0.22 | 1.3933 | 1.1 | 0.76 | 0.25333 | 0.63333 | 1.6467 | 1.3
381 | 20 | 0.12 | 127 | 0.1 | 0.22 | 1.397 | 1.1 | 0.762 | 0.254 | 0.635 | 1.651 | 1.3
382 | 20 | 0.12 | 127.3333 | 0.1 | 0.22 | 1.4007 | 1.1 | 0.764 | 0.25467 | 0.63667 | 1.6553 | 1.3
383 | 20 | 0.12 | 127.6667 | 0.1 | 0.22 | 1.4043 | 1.1 | 0.766 | 0.25533 | 0.63833 | 1.6597 | 1.3
384 | 20 | 0.12 | 128 | 0.1 | 0.22 | 1.408 | 1.1 | 0.768 | 0.256 | 0.64 | 1.664 | 1.3
385 | 20 | 0.12 | 128.3333 | 0.1 | 0.22 | 1.4117 | 1.1 | 0.77 | 0.25667 | 0.64167 | 1.6683 | 1.3
386 | 20 | 0.12 | 128.6667 | 0.1 | 0.22 | 1.4153 | 1.1 | 0.772 | 0.25733 | 0.64333 | 1.6727 | 1.3
387 | 20 | 0.12 | 129 | 0.1 | 0.22 | 1.419 | 1.1 | 0.774 | 0.258 | 0.645 | 1.677 | 1.3
388 | 20 | 0.12 | 129.3333 | 0.1 | 0.22 | 1.4227 | 1.1 | 0.776 | 0.25867 | 0.64667 | 1.6813 | 1.3
389 | 20 | 0.12 | 129.6667 | 0.1 | 0.22 | 1.4263 | 1.1 | 0.778 | 0.25933 | 0.64833 | 1.6857 | 1.3
390 | 20 | 0.12 | 130 | 0.1 | 0.22 | 1.43 | 1.1 | 0.78 | 0.26 | 0.65 | 1.69 | 1.3
391 | 20 | 0.12 | 130.3333 | 0.1 | 0.22 | 1.4337 | 1.1 | 0.782 | 0.26067 | 0.65167 | 1.6943 | 1.3
392 | 20 | 0.12 | 130.6667 | 0.1 | 0.22 | 1.4373 | 1.1 | 0.784 | 0.26133 | 0.65333 | 1.6987 | 1.3
393 | 20 | 0.12 | 131 | 0.1 | 0.22 | 1.441 | 1.1 | 0.786 | 0.262 | 0.655 | 1.703 | 1.3
394 | 20 | 0.12 | 131.3333 | 0.1 | 0.22 | 1.4447 | 1.1 | 0.788 | 0.26267 | 0.65667 | 1.7073 | 1.3
395 | 20 | 0.12 | 131.6667 | 0.1 | 0.22 | 1.4483 | 1.1 | 0.79 | 0.26333 | 0.65833 | 1.7117 | 1.3
396 | 20 | 0.12 | 132 | 0.1 | 0.22 | 1.452 | 1.1 | 0.792 | 0.264 | 0.66 | 1.716 | 1.3
397 | 20 | 0.12 | 132.3333 | 0.1 | 0.22 | 1.4557 | 1.1 | 0.794 | 0.26467 | 0.66167 | 1.7203 | 1.3
398 | 20 | 0.12 | 132.6667 | 0.1 | 0.22 | 1.4593 | 1.1 | 0.796 | 0.26533 | 0.66333 | 1.7247 | 1.3
399 | 20 | 0.12 | 133 | 0.1 | 0.22 | 1.463 | 1.1 | 0.798 | 0.266 | 0.665 | 1.729 | 1.3
400 | 20 | 0.12 | 133.3333 | 0.1 | 0.22 | 1.4667 | 1.1 | 0.8 | 0.26667 | 0.66667 | 1.7333 | 1.3
401 | 20 | 0.12 | 133.6667 | 0.1 | 0.22 | 1.4703 | 1.1 | 0.802 | 0.26733 | 0.66833 | 1.7377 | 1.3
402 | 20 | 0.12 | 134 | 0.1 | 0.22 | 1.474 | 1.1 | 0.804 | 0.268 | 0.67 | 1.742 | 1.3
403 | 20 | 0.12 | 134.3333 | 0.1 | 0.22 | 1.4777 | 1.1 | 0.806 | 0.26867 | 0.67167 | 1.7463 | 1.3
404 | 20 | 0.12 | 134.6667 | 0.1 | 0.22 | 1.4813 | 1.1 | 0.808 | 0.26933 | 0.67333 | 1.7507 | 1.3
405 | 20 | 0.12 | 135 | 0.1 | 0.22 | 1.485 | 1.1 | 0.81 | 0.27 | 0.675 | 1.755 | 1.3
406 | 20 | 0.12 | 135.3333 | 0.1 | 0.22 | 1.4887 | 1.1 | 0.812 | 0.27067 | 0.67667 | 1.7593 | 1.3
407 | 20 | 0.12 | 135.6667 | 0.1 | 0.22 | 1.4923 | 1.1 | 0.814 | 0.27133 | 0.67833 | 1.7637 | 1.3
408 | 20 | 0.12 | 136 | 0.1 | 0.22 | 1.496 | 1.1 | 0.816 | 0.272 | 0.68 | 1.768 | 1.3
409 | 20 | 0.12 | 136.3333 | 0.1 | 0.22 | 1.4997 | 1.1 | 0.818 | 0.27267 | 0.68167 | 1.7723 | 1.3
410 | 20 | 0.12 | 136.6667 | 0.1 | 0.22 | 1.5033 | 1.1 | 0.82 | 0.27333 | 0.68333 | 1.7767 | 1.3
411 | 20 | 0.12 | 137 | 0.1 | 0.22 | 1.507 | 1.1 | 0.822 | 0.274 | 0.685 | 1.781 | 1.3
412 | 20 | 0.12 | 137.3333 | 0.1 | 0.22 | 1.5107 | 1.1 | 0.824 | 0.27467 | 0.68667 | 1.7853 | 1.3
413 | 20 | 0.12 | 137.6667 | 0.1 | 0.22 | 1.5143 | 1.1 | 0.826 | 0.27533 | 0.68833 | 1.7897 | 1.3
414 | 20 | 0.12 | 138 | 0.1 | 0.22 | 1.518 | 1.1 | 0.828 | 0.276 | 0.69 | 1.794 | 1.3
415 | 20 | 0.12 | 138.3333 | 0.1 | 0.22 | 1.5217 | 1.1 | 0.83 | 0.27667 | 0.69167 | 1.7983 | 1.3
416 | 20 | 0.12 | 138.6667 | 0.1 | 0.22 | 1.5253 | 1.1 | 0.832 | 0.27733 | 0.69333 | 1.8027 | 1.3
417 | 20 | 0.12 | 139 | 0.1 | 0.22 | 1.529 | 1.1 | 0.834 | 0.278 | 0.695 | 1.807 | 1.3
418 | 20 | 0.12 | 139.3333 | 0.1 | 0.22 | 1.5327 | 1.1 | 0.836 | 0.27867 | 0.69667 | 1.8113 | 1.3
419 | 20 | 0.12 | 139.6667 | 0.1 | 0.22 | 1.5363 | 1.1 | 0.838 | 0.27933 | 0.69833 | 1.8157 | 1.3
420 | 20 | 0.12 | 140 | 0.1 | 0.22 | 1.54 | 1.1 | 0.84 | 0.28 | 0.7 | 1.82 | 1.3
421 | 20 | 0.12 | 140.3333 | 0.1 | 0.22 | 1.5437 | 1.1 | 0.842 | 0.28067 | 0.70167 | 1.8243 | 1.3
422 | 20 | 0.12 | 140.6667 | 0.1 | 0.22 | 1.5473 | 1.1 | 0.844 | 0.28133 | 0.70333 | 1.8287 | 1.3
423 | 20 | 0.12 | 141 | 0.1 | 0.22 | 1.551 | 1.1 | 0.846 | 0.282 | 0.705 | 1.833 | 1.3
424 | 20 | 0.12 | 141.3333 | 0.1 | 0.22 | 1.5547 | 1.1 | 0.848 | 0.28267 | 0.70667 | 1.8373 | 1.3
425 | 20 | 0.12 | 141.6667 | 0.1 | 0.22 | 1.5583 | 1.1 | 0.85 | 0.28333 | 0.70833 | 1.8417 | 1.3
426 | 20 | 0.12 | 142 | 0.1 | 0.22 | 1.562 | 1.1 | 0.852 | 0.284 | 0.71 | 1.846 | 1.3
427 | 20 | 0.12 | 142.3333 | 0.1 | 0.22 | 1.5657 | 1.1 | 0.854 | 0.28467 | 0.71167 | 1.8503 | 1.3
428 | 20 | 0.12 | 142.6667 | 0.1 | 0.22 | 1.5693 | 1.1 | 0.856 | 0.28533 | 0.71333 | 1.8547 | 1.3
429 | 20 | 0.12 | 143 | 0.1 | 0.22 | 1.573 | 1.1 | 0.858 | 0.286 | 0.715 | 1.859 | 1.3
430 | 20 | 0.12 | 143.3333 | 0.1 | 0.22 | 1.5767 | 1.1 | 0.86 | 0.28667 | 0.71667 | 1.8633 | 1.3
431 | 20 | 0.12 | 143.6667 | 0.1 | 0.22 | 1.5803 | 1.1 | 0.862 | 0.28733 | 0.71833 | 1.8677 | 1.3
432 | 20 | 0.12 | 144 | 0.1 | 0.22 | 1.584 | 1.1 | 0.864 | 0.288 | 0.72 | 1.872 | 1.3
433 | 20 | 0.12 | 144.3333 | 0.1 | 0.22 | 1.5877 | 1.1 | 0.866 | 0.28867 | 0.72167 | 1.8763 | 1.3
434 | 20 | 0.12 | 144.6667 | 0.1 | 0.22 | 1.5913 | 1.1 | 0.868 | 0.28933 | 0.72333 | 1.8807 | 1.3
435 | 20 | 0.12 | 145 | 0.1 | 0.22 | 1.595 | 1.1 | 0.87 | 0.29 | 0.725 | 1.885 | 1.3
436 | 20 | 0.12 | 145.3333 | 0.1 | 0.22 | 1.5987 | 1.1 | 0.872 | 0.29067 | 0.72667 | 1.8893 | 1.3
437 | 20 | 0.12 | 145.6667 | 0.1 | 0.22 | 1.6023 | 1.1 | 0.874 | 0.29133 | 0.72833 | 1.8937 | 1.3
438 | 20 | 0.12 | 146 | 0.1 | 0.22 | 1.606 | 1.1 | 0.876 | 0.292 | 0.73 | 1.898 | 1.3
439 | 20 | 0.12 | 146.3333 | 0.1 | 0.22 | 1.6097 | 1.1 | 0.878 | 0.29267 | 0.73167 | 1.9023 | 1.3
440 | 20 | 0.12 | 146.6667 | 0.1 | 0.22 | 1.6133 | 1.1 | 0.88 | 0.29333 | 0.73333 | 1.9067 | 1.3
441 | 20 | 0.12 | 147 | 0.1 | 0.22 | 1.617 | 1.1 | 0.882 | 0.294 | 0.735 | 1.911 | 1.3
442 | 20 | 0.12 | 147.3333 | 0.1 | 0.22 | 1.6207 | 1.1 | 0.884 | 0.29467 | 0.73667 | 1.9153 | 1.3
443 | 20 | 0.12 | 147.6667 | 0.1 | 0.22 | 1.6243 | 1.1 | 0.886 | 0.29533 | 0.73833 | 1.9197 | 1.3
444 | 20 | 0.12 | 148 | 0.1 | 0.22 | 1.628 | 1.1 | 0.888 | 0.296 | 0.74 | 1.924 | 1.3
445 | 20 | 0.12 | 148.3333 | 0.1 | 0.22 | 1.6317 | 1.1 | 0.89 | 0.29667 | 0.74167 | 1.9283 | 1.3
446 | 20 | 0.12 | 148.6667 | 0.1 | 0.22 | 1.6353 | 1.1 | 0.892 | 0.29733 | 0.74333 | 1.9327 | 1.3
447 | 20 | 0.12 | 149 | 0.1 | 0.22 | 1.639 | 1.1 | 0.894 | 0.298 | 0.745 | 1.937 | 1.3
448 | 20 | 0.12 | 149.3333 | 0.1 | 0.22 | 1.6427 | 1.1 | 0.896 | 0.29867 | 0.74667 | 1.9413 | 1.3
449 | 20 | 0.12 | 149.6667 | 0.1 | 0.22 | 1.6463 | 1.1 | 0.898 | 0.29933 | 0.74833 | 1.9457 | 1.3
450 | 20 | 0.12 | 150 | 0.1 | 0.22 | 1.65 | 1.1 | 0.9 | 0.3 | 0.75 | 1.95 | 1.3
451 | 20 | 0.12 | 150.3333 | 0.1 | 0.22 | 1.6537 | 1.1 | 0.902 | 0.30067 | 0.75167 | 1.9543 | 1.3
452 | 20 | 0.12 | 150.6667 | 0.1 | 0.22 | 1.6573 | 1.1 | 0.904 | 0.30133 | 0.75333 | 1.9587 | 1.3
453 | 20 | 0.12 | 151 | 0.1 | 0.22 | 1.661 | 1.1 | 0.906 | 0.302 | 0.755 | 1.963 | 1.3
454 | 20 | 0.12 | 151.3333 | 0.1 | 0.22 | 1.6647 | 1.1 | 0.908 | 0.30267 | 0.75667 | 1.9673 | 1.3
455 | 20 | 0.12 | 151.6667 | 0.1 | 0.22 | 1.6683 | 1.1 | 0.91 | 0.30333 | 0.75833 | 1.9717 | 1.3
456 | 20 | 0.12 | 152 | 0.1 | 0.22 | 1.672 | 1.1 | 0.912 | 0.304 | 0.76 | 1.976 | 1.3
457 | 20 | 0.12 | 152.3333 | 0.1 | 0.22 | 1.6757 | 1.1 | 0.914 | 0.30467 | 0.76167 | 1.9803 | 1.3
458 | 20 | 0.12 | 152.6667 | 0.1 | 0.22 | 1.6793 | 1.1 | 0.916 | 0.30533 | 0.76333 | 1.9847 | 1.3
459 | 20 | 0.12 | 153 | 0.1 | 0.22 | 1.683 | 1.1 | 0.918 | 0.306 | 0.765 | 1.989 | 1.3
460 | 20 | 0.12 | 153.3333 | 0.1 | 0.22 | 1.6867 | 1.1 | 0.92 | 0.30667 | 0.76667 | 1.9933 | 1.3
461 | 20 | 0.12 | 153.6667 | 0.1 | 0.22 | 1.6903 | 1.1 | 0.922 | 0.30733 | 0.76833 | 1.9977 | 1.3
462 | 20 | 0.12 | 154 | 0.1 | 0.22 | 1.694 | 1.1 | 0.924 | 0.308 | 0.77 | 2.002 | 1.3
463 | 20 | 0.12 | 154.3333 | 0.1 | 0.22 | 1.6977 | 1.1 | 0.926 | 0.30867 | 0.77167 | 2.0063 | 1.3
464 | 20 | 0.12 | 154.6667 | 0.1 | 0.22 | 1.7013 | 1.1 | 0.928 | 0.30933 | 0.77333 | 2.0107 | 1.3
465 | 20 | 0.12 | 155 | 0.1 | 0.22 | 1.705 | 1.1 | 0.93 | 0.31 | 0.775 | 2.015 | 1.3
466 | 20 | 0.12 | 155.3333 | 0.1 | 0.22 | 1.7087 | 1.1 | 0.932 | 0.31067 | 0.77667 | 2.0193 | 1.3
467 | 20 | 0.12 | 155.6667 | 0.1 | 0.22 | 1.7123 | 1.1 | 0.934 | 0.31133 | 0.77833 | 2.0237 | 1.3
468 | 20 | 0.12 | 156 | 0.1 | 0.22 | 1.716 | 1.1 | 0.936 | 0.312 | 0.78 | 2.028 | 1.3
469 | 20 | 0.12 | 156.3333 | 0.1 | 0.22 | 1.7197 | 1.1 | 0.938 | 0.31267 | 0.78167 | 2.0323 | 1.3
470 | 20 | 0.12 | 156.6667 | 0.1 | 0.22 | 1.7233 | 1.1 | 0.94 | 0.31333 | 0.78333 | 2.0367 | 1.3
471 | 20 | 0.12 | 157 | 0.1 | 0.22 | 1.727 | 1.1 | 0.942 | 0.314 | 0.785 | 2.041 | 1.3
472 | 20 | 0.12 | 157.3333 | 0.1 | 0.22 | 1.7307 | 1.1 | 0.944 | 0.31467 | 0.78667 | 2.0453 | 1.3
473 | 20 | 0.12 | 157.6667 | 0.1 | 0.22 | 1.7343 | 1.1 | 0.946 | 0.31533 | 0.78833 | 2.0497 | 1.3
474 | 20 | 0.12 | 158 | 0.1 | 0.22 | 1.738 | 1.1 | 0.948 | 0.316 | 0.79 | 2.054 | 1.3
475 | 20 | 0.12 | 158.3333 | 0.1 | 0.22 | 1.7417 | 1.1 | 0.95 | 0.31667 | 0.79167 | 2.0583 | 1.3
476 | 20 | 0.12 | 158.6667 | 0.1 | 0.22 | 1.7453 | 1.1 | 0.952 | 0.31733 | 0.79333 | 2.0627 | 1.3
477 | 20 | 0.12 | 159 | 0.1 | 0.22 | 1.749 | 1.1 | 0.954 | 0.318 | 0.795 | 2.067 | 1.3
478 | 20 | 0.12 | 159.3333 | 0.1 | 0.22 | 1.7527 | 1.1 | 0.956 | 0.31867 | 0.79667 | 2.0713 | 1.3
479 | 20 | 0.12 | 159.6667 | 0.1 | 0.22 | 1.7563 | 1.1 | 0.958 | 0.31933 | 0.79833 | 2.0757 | 1.3
480 | 20 | 0.12 | 160 | 0.1 | 0.22 | 1.76 | 1.1 | 0.96 | 0.32 | 0.8 | 2.08 | 1.3
481 | 20 | 0.12 | 160.3333 | 0.1 | 0.22 | 1.7637 | 1.1 | 0.962 | 0.32067 | 0.80167 | 2.0843 | 1.3
482 | 20 | 0.12 | 160.6667 | 0.1 | 0.22 | 1.7673 | 1.1 | 0.964 | 0.32133 | 0.80333 | 2.0887 | 1.3
483 | 20 | 0.12 | 161 | 0.1 | 0.22 | 1.771 | 1.1 | 0.966 | 0.322 | 0.805 | 2.093 | 1.3
484 | 20 | 0.12 | 161.3333 | 0.1 | 0.22 | 1.7747 | 1.1 | 0.968 | 0.32267 | 0.80667 | 2.0973 | 1.3
485 | 20 | 0.12 | 161.6667 | 0.1 | 0.22 | 1.7783 | 1.1 | 0.97 | 0.32333 | 0.80833 | 2.1017 | 1.3
486 | 20 | 0.12 | 162 | 0.1 | 0.22 | 1.782 | 1.1 | 0.972 | 0.324 | 0.81 | 2.106 | 1.3
487 | 20 | 0.12 | 162.3333 | 0.1 | 0.22 | 1.7857 | 1.1 | 0.974 | 0.32467 | 0.81167 | 2.1103 | 1.3
488 | 20 | 0.12 | 162.6667 | 0.1 | 0.22 | 1.7893 | 1.1 | 0.976 | 0.32533 | 0.81333 | 2.1147 | 1.3
489 | 20 | 0.12 | 163 | 0.1 | 0.22 | 1.793 | 1.1 | 0.978 | 0.326 | 0.815 | 2.119 | 1.3
490 | 20 | 0.12 | 163.3333 | 0.1 | 0.22 | 1.7967 | 1.1 | 0.98 | 0.32667 | 0.81667 | 2.1233 | 1.3
491 | 20 | 0.12 | 163.6667 | 0.1 | 0.22 | 1.8003 | 1.1 | 0.982 | 0.32733 | 0.81833 | 2.1277 | 1.3
492 | 20 | 0.12 | 164 | 0.1 | 0.22 | 1.804 | 1.1 | 0.984 | 0.328 | 0.82 | 2.132 | 1.3
493 | 20 | 0.12 | 164.3333 | 0.1 | 0.22 | 1.8077 | 1.1 | 0.986 | 0.32867 | 0.82167 | 2.1363 | 1.3
494 | 20 | 0.12 | 164.6667 | 0.1 | 0.22 | 1.8113 | 1.1 | 0.988 | 0.32933 | 0.82333 | 2.1407 | 1.3
495 | 20 | 0.12 | 165 | 0.1 | 0.22 | 1.815 | 1.1 | 0.99 | 0.33 | 0.825 | 2.145 | 1.3
496 | 20 | 0.12 | 165.3333 | 0.1 | 0.22 | 1.8187 | 1.1 | 0.992 | 0.33067 | 0.82667 | 2.1493 | 1.3
497 | 20 | 0.12 | 165.6667 | 0.1 | 0.22 | 1.8223 | 1.1 | 0.994 | 0.33133 | 0.82833 | 2.1537 | 1.3
498 | 20 | 0.12 | 166 | 0.1 | 0.22 | 1.826 | 1.1 | 0.996 | 0.332 | 0.83 | 2.158 | 1.3
499 | 20 | 0.12 | 166.3333 | 0.1 | 0.22 | 1.8297 | 1.1 | 0.998 | 0.33267 | 0.83167 | 2.1623 | 1.3
500 | 20 | 0.12 | 166.6667 | 0.1 | 0.22 | 1.8333 | 1.1 | 1 | 0.33333 | 0.83333 | 2.1667 | 1.3
501 | 20 | 0.12 | 167 | 0.1 | 0.22 | 1.837 | 1.1 | 1.002 | 0.334 | 0.835 | 2.171 | 1.3
502 | 20 | 0.12 | 167.3333 | 0.1 | 0.22 | 1.8407 | 1.1 | 1.004 | 0.33467 | 0.83667 | 2.1753 | 1.3
503 | 20 | 0.12 | 167.6667 | 0.1 | 0.22 | 1.8443 | 1.1 | 1.006 | 0.33533 | 0.83833 | 2.1797 | 1.3
504 | 20 | 0.12 | 168 | 0.1 | 0.22 | 1.848 | 1.1 | 1.008 | 0.336 | 0.84 | 2.184 | 1.3
505 | 20 | 0.12 | 168.3333 | 0.1 | 0.22 | 1.8517 | 1.1 | 1.01 | 0.33667 | 0.84167 | 2.1883 | 1.3
506 | 20 | 0.12 | 168.6667 | 0.1 | 0.22 | 1.8553 | 1.1 | 1.012 | 0.33733 | 0.84333 | 2.1927 | 1.3
507 | 20 | 0.12 | 169 | 0.1 | 0.22 | 1.859 | 1.1 | 1.014 | 0.338 | 0.845 | 2.197 | 1.3
508 | 20 | 0.12 | 169.3333 | 0.1 | 0.22 | 1.8627 | 1.1 | 1.016 | 0.33867 | 0.84667 | 2.2013 | 1.3
509 | 20 | 0.12 | 169.6667 | 0.1 | 0.22 | 1.8663 | 1.1 | 1.018 | 0.33933 | 0.84833 | 2.2057 | 1.3
510 | 20 | 0.12 | 170 | 0.1 | 0.22 | 1.87 | 1.1 | 1.02 | 0.34 | 0.85 | 2.21 | 1.3 | %FS error at 1 SLPM | %RD error ar 1 SLPM | Tot added uncertainty (%RD) | Upper limit of totalizer uncertainty if flowing 1 SLPM (%RD)
511 | 20 | 0.12 | 170.3333 | 0.1 | 0.22 | 1.8737 | 1.1 | 1.022 | 0.34067 | 0.85167 | 2.2143 | 1.3 | 0.04 | 4 | 0.5 | 4.5